## Slide 1: Scanning Electron Microscopy
Nanotechnology Lab at North Seattle College
Notes: 45 minutes

## Slide 2: Microscope Background (LM vs EM)
Light Microscopes (LM)
  Visible light (700-400nm) to illuminate the sample
  Limited to ~1500X magnification, ~200nm resolution
Electron Microscopes (EM)
  Focused beam of electrons  shorter wavelength  better resolution
  Transmission Electron Microscope (TEM)
  Scanning Electron Microscope (SEM)
    ~10X – 300,000X magnification, <10nm resolution
Notes: Microscope History
Optical Microscope - ~1600s Galileo
Early 1920s – discovered electrons could be focused by magnetic field
Resolution at high magnification. Resolution can be defined as the least distance between two closely opposed points, at which they may be recognized as two separate entities. The best resolution possible in a LM is about 200 nm whereas a typical SEM has a resolution of better than 10 nm (typically 5 nm).
Depth of field i.e. the height of a specimen that appears in focus in an image - more than 300 times the depth of field compared to the LM. This means that great topographical detail can be obtained. For many users, the three dimensional (3D) appearance of the specimen image, is the most valuable feature of the SEM. This is because such images, even at low magnifications, can provide much more information about a specimen than is available using the LM. The use of "stereo pair" SEM images can give even greater information about the sample.
Microanalysis i.e. the analysis of sample composition including information about chemical composition, as well as crystallographic, magnetic and electrical characteristics.
One drawback to the use of the SEM is that it operates under vacuum and in many SEMs the samples must be rendered conductive to be viewed. This is often achieved by coating with a very thin layer of metal or carbon
Other Notes:
SEM cannot take color images. Color can be added artificially in SEM images. Some SEMs can collect true color images using a wavelength selective cathodolumenence (CL) detector.
EM cannot image through water. Note: An ESEM using a wet Scanning Transmission Electron Microscope (STEM) detector can be used to image through thin water films.
The SEM cannot reliably image charged molecules that are mobile in a matrix. For example, some species (e.g. Na+) are volatile under the electron beam because the negative electron beam exerts a force on charged material.
SEM is not good for quantifying surface roughness at small scale. Atomic Force Microscope (Scanning Probe Microscopy) is more useful for this task.
Measurements involving height (z-axis) cannot be taken directly in an SEM. This requires two images that have been tilted relative to one another to create a 3D image, and specialised processing software.
Generally, SEMs are not used for experiments involving liquids, chemical reactions, and air-gas systems although some specialized machines and sample chambers do allow for these experiments.
The resolution of the SEM is not high enough to image individual atoms (use a transmission electron microscope).
Elemental analysis below micron scale.
Note: Analysis in the < 7kV range can provide elemental information on the sub-micron scale but is often problematical.

## Slide 3: Microscope Resolution (LM vs EM)
Abbe’s Equation
Electron Microscopes (EM):
nm

## Slide 4: Light Microscope in Action
Load your sample on the stage
Adjust lenses to bring image into focus
In-focus image of student produced MEMS.
Notes: MEMS: microelectronic and microelectromechanical systems – used in chemical/biochemical sensors
For example: An accelerometer, which is used when you play Nintendo wii to let the device know when you turn or move direction/angle/speed, sensors

## Slide 5
M. von Ardenne's first SEM
Scanning Electron Microscope. In Wikipedia. Retrieved June 10, 2015, from https://en.wikipedia.org/wiki/Scanning_electron_microscope
Aspex Explorer SEM at NSC
Notes: "Ernst Ruska Electron Microscope - Deutsches Museum - Munich-edit" by J Brew, uploaded on the English-speaking Wikipedia by en:User:Hat'nCoat. - originally posted to Flickr as Electron Microscope Deutsches Museum.
Licensed under CC BY-SA 3.0 via Wikimedia Commons – http://commons.wikimedia.org/wiki/File:Ernst_Ruska_Electron_Microscope_-_Deutsches_Museum_-_Munich-edit.jpg#/media/File:Ernst_Ruska_Electron_Microscope_-_Deutsches_Museum_-_Munich-edit.jpg
"First Scanning Electron Microscope with high resolution from Manfred von Ardenne 1937“
by Rechteinhaber: Dr. rer. nat. Alexander von Ardenne - by inheritance from my father Prof. Dr. h. c. mult. Manfred von Ardenne; he died in 1997 (online at: www.ardenne.de/med).
Licensed under CC BY-SA 3.0 via Wikimedia Commons - http://commons.wikimedia.org/wiki/File:First_Scanning_Electron_Microscope_with_high_resolution_from_Manfred_von_Ardenne_1937.jpg#/media/File:First_Scanning_Electron_Microscope_with_high_resolution_from_Manfred_von_Ardenne_1937.jpg

## Slide 6: Scanning Electron Microscope in Action

## Slide 7: Electron Gun
Thermionic Tungsten (W) Filament
  Current heats up filament
  Photons and electrons emitted
Wehnelt Cap
  small hole only allowing electrons emission to tip
Accelerating Anode
  Attracts electrons through center and EM field propels electrons forward
http://www.ammrf.org.au/myscope/sem
Notes: Wehnelt cap: smaller hole for improved characteristics
Anode (+) accelerates the electrons by attracting them throgh the center – entering magnetic field allows them to keep propelling forward
Filament current (Load Current) : the current that is passed through the filament in the electron gun to produce the electrons for the beam.
http://www4.nau.edu/microanalysis/microprobe/Column-ElectronGun.html

## Slide 8: Scanning Electron Microscope
Electron beam generated
Electromagnetic Condenser Lenses
  focus beam as it leaves the filament’
Deflection (Scanning) Coils
  Moves beam across sample
Objective Lens (Final) Lens
  Focuses beam near sample surface
http://www.ammrf.org.au/myscope/sem
Notes: Electron Column: Focuses the beam of electrons that is generated by the electron gun
Condenser Lenses: An electromagnetic lens located just below the electron gun to focus the electron beam emitted by the filament.
Objective Lens (Aperture) : The lens nearest to the specimen and used to focus the electron beam onto the specimen surface; this aperture strip of metal that is placed in the path of the electron beam in order to restrict or limit electron progress down the machine column. An aperture stops off-axis or off-energy electrons.
Scanning coils (Copper Coils – iron magnet incasing): Beam is deflected in the x and y axes by a pair of scanning coils (or deflector plates). These are found below the condenser lens system and associated with the objective lens area. The deflection enables the movement of the beam, as lines of consecutive dwell-times for the beam spot, across a rectangular area of the sample surface. This is also called rastering.
Stigmator is used to make the probe circular. The sigmator is comprised of electromagnetic coils placed around in in quadruple, sextuple or octagonal orientations.
Accelerating voltage (kV or keV) is the voltage difference between the filament and the anode which accelerates the electron beam towards the anode.

## Slide 9: SEM: Electron Beam Interactions
Secondary Electrons
  Inelastic scattering
  Ejection of loosely (valence) bound electrons
  Low energy ~2-50eV e- collected by detector
  Important for topographical characterization
Secondary Electron Detector Images
  Sample areas faced away or blocked - darker
+
+
+
+
+
+
valence e-
core e-
nucleus
Notes: SE: Produced through inelastic scattering that results in the ejection of loosely bound electrons from the specimen. Secondary electrons have energies from ~2-50eV.
Scintillator: A material that emits light when excited by electromagnetic radiation or a beam of subatomic particles, e.g. electrons.
X-ray wavelength: 0.01nm – 10nm

## Slide 10: SEM: Electron Beam Interactions
Backscattered Electrons
  Elastic scattering
  Reflection – electrons bounce back
  High energy >50eV collected
  Qualitative chemical analysis
Backscatter Electron Detector Images
  Higher atomic weight 
  larger nuclei  brighter
+
+
+
+
+
+
valence e-
core e-
nucleus
http://www.ammrf.org.au/myscope/sem

## Slide 11: SEM: Electron Beam Interactions (Cont’d)
Core Electron Ejection
  Tightly bound electrons removed leaves atom in excited state
  Higher energy level electron transition to fill hole
    Emits x-ray photon or causes Auger electron emitted
Heat and Light

## Slide 12: SEM Sample Conditions
Samples under high vacuum
  ~10-5 – 10-7 torr (atmospheric pressure ~760 torr)
Dry
Sample is conductive or coated with thin layer of conductive material
  Gold (Au), Palladium/Gold Alloy (Pd/Au), Carbon (C)
Notes: VP/LV (Variable Pressure, Low Voltage) SEM – for sensitive, non-conductive, uncoated samples. Must use BSED

## Slide 13: Energy Dispersive X-ray Spectroscopy
http://www.ammrf.org.au/myscope/sem

## Slide 14: SEM Uses
Materials Science
  Failure Analysis
  Quality Control
Forensic Science
  Examine evidence, chemical analysis
Biological and Medicinal Science
  Surface characteristics
  Identification
Archeology and Museum Science
Geological Samples
  Chemical Composition (BSED and EDX)
  Morphology
Art

## Slide 15: AFM Tip and Caterpillar

## Slide 16: Cork and Memory Foam

## Slide 17: Yarrow Leaf and Glitter

## Slide 18: Coffee Bean and Ball on Sponge

## Slide 19: Fungus and Bug in Web

## Slide 20: Tooth and Ant Eye