## Slide 1
ACADs (08-006) Covered
Keywords
Errors, accuracy, count rate, background, count time, equipment efficiency, sample volume, sample geometry, moisture absorption, standard deviation, distribution, confidence, chi square.
Description
Supporting Material
Counting Statistics Instructor Notes
1.1.1.2 | 3.2.3.19 | 3.2.3.20 | 4.13.1 | 4.13.2 | 4.13.4 |  | 
Notes: <number>

## Slide 2
HPT001.011
Revision 3
Page of 45
TP-<number>
Counting Statistics
31
Notes: <number>

## Slide 3: Statistical Accuracy
Factors that affect the statistical accuracy of Radioactivity Measurements:
  Count rate
  Background
  Count time
  Equipment Efficiency
  Sample Volume
  Sample geometry
  Moisture Absorption
HPT001.011
Revision 3
Page of 45
TP-<number>
32

## Slide 4: Statistical Accuracy
Errors Relating to Arithmetic Calculations:
  Human errors
  Variations in individual measurements
HPT001.011
Revision 3
Page of 45
TP-<number>
Notes:   Human errors such as incorrect mathematical manipulations or incorrect decimal positions.
  Errors due to variations in individual measurements carried and increased by the mathematical manipulations.
<number>

## Slide 5: Statistical Accuracy
Errors Relating to Radioactivity Measurements
  Random disintegration for radioactive atoms.
  Random emission of radiations when an atom decays.
  Detector related errors
  Radiation Measurement Technique Errors
HPT001.011
Revision 3
Page of 45
TP-<number>
Notes: 9. Errors Relating to Radioactivity Measurements
a. Random disintegration for radioactive atoms.
b. Random emission of radiations when an atom decays.
c. Detector related errors
(1) Every radiation emitted will not reach the detector.
(2) Every radiation entering the detector will not be counted.
(a) This is especially true at high counting rates.
(b) A certain amount of time is needed for the electronics to recover from one count before it can count another disintegration event.
(3) There will be electronic and detector variations due to changes in temperature and applied voltage.
(4) In detectors that use counting gases, the sealed gases break down gradually, and the flow of non-sealed gases can vary.
d. Radiation Measurement Technique Errors
(1) The absorption of beta radiation by windows or walls of the detector, and by a cover over the source, and even by air reduces the number of beta particles reaching the detector.
(2) When gamma rays are absorbed by materials outside the detector, secondary radiations are produced that can enter the detector.
(3) Self-absorption of the radiation by other atoms in the source.
(4) Scattering and backscattering can occur when radiation interacts with matter both in and out of the detector with part, but not all of the energy being deposited in the detector, i.e., part escapes and is not detected or recorded.
<number>

## Slide 6: Error Reduction
Peer Check
STAR
Procedure Use and Adherence
HPT001.011
Revision 3
Page of 45
TP-<number>
33
Events

## Slide 7: Accuracy and Precision
HPT001.011
Revision 3
Page of 45
TP-<number>
34

## Slide 8: Standard Deviation
Represented by the Greek symbol sigma 
One  is the distance from the peak out to a vertical line enclosing 34.15% of the total area under the curve
HPT001.011
Revision 3
Page of 45
TP-<number>
35
Where’s the mean?

## Slide 9: Frequency Distribution
HPT001.011
Revision 3
Page of 45
TP-<number>
36
Data is plotted on a histogram
Height of bar represents frequency of occurrence

## Slide 10: Poisson Distribution
HPT001.011
Revision 3
Page of 45
TP-<number>
37
Probability of “success” is low
Number of trials is high

## Slide 11: Gaussian Distribution
HPT001.011
Revision 3
Page of 45
TP-<number>
38
Symmetrical about the mean
One  includes 68.3% of area under curve

## Slide 12: Confidence Level
HPT001.011
Revision 3
Page of 45
TP-<number>
39
2 = 95.4 % confidence level
1 = 68.3 % confidence level
Notes: <number>

## Slide 13: Minimum Detectable Count Rate
Calculated using the equation:
HPT001.011
Revision 3
Page of 45
TP-<number>
40
MDC = 2.71 + 3.3 [B(tb+ts)/tb]1/2
where: MDC is the minimum dectectable counts;
B is the background counts;
tb is the background counting time, minutes;
ts is the sample counting time, minutes.
MDCR = MDC/ts

## Slide 14: MDCR Application
If gross count rate is > (Bkd + MDCR):
HPT001.011
Revision 3
Page of 45
TP-<number>
41
It may be concluded with 95% confidence that radioactivity is present above natural background. Calculate results using normal processes.

## Slide 15: MDCR Application (cont’d)
If gross count rate is < (Bkd + MDCR):
HPT001.011
Revision 3
Page of 45
TP-<number>
42
record as
"< MDA".

## Slide 16: Lower Limit of Detection
Calculated using the equation:
HPT001.011
Revision 3
Page of 45
TP-<number>
43
LLD (Ci/cc) = 4.66 b
(2.22E6)(E)(V)(Y)(D)
where: V is the sample volume in cc;
E is the counter efficiency (cts/dis);
Y is the chemical yield if applicable;
D is the decay correction for delayed count on sample.
2.22E6 is a conversion factor - dpm per Ci

## Slide 17: Chi-Square Test
Calculated using the equation:
HPT001.011
Revision 3
Page of 45
TP-<number>
44
x2 = (n-)2

where: n = the data for each count;
 = the average of the individual counts;
 = n/N
N = the number of observations (usually 21)

## Slide 18: Chi-Square Test
Counts (n) | (n-) | (n-)2
52377 | +116 | 13456
52202 | -59 | 3481
52102 | -159 | 25281
52385 | +124 | 15376
52427 | +166 | 27556
52133 | -128 | 16384
52460 | +199 | 39601
52040 | -221 | 48841
52188 | -73 | 5329
52003 | -258 | 66564
52680 | +419 | 175561
51972 | -289 | 83521
52441 | +180 | 32400
52153 | -108 | 11664
52388 | +127 | 16129
52287 | +26 | 676
52108 | -153 | 23409
52399 | +138 | 19044
52290 | +29 | 841
52185 | -76 | 5776
HPT001.011
Revision 3
Page of 45
TP-<number>
45