## Slide 1
ACADs (08-006) Covered
Keywords
Efficiency, LLD, CPM, DPM, relative efficiency, absolute efficiency, standard deviation, confidence, count time.
Description
Supporting Material
Detector Efficiency
1.1.1.2 | 3.2.3.20 | 3.2.3.21 | 3.2.3.22 | 4.13.3 |  |  | 

## Slide 2: Detector Efficiency
IRAD 2371
Week 3

## Slide 3: Efficiency
Very few detectors will count every interaction
Each detector will have its own counting efficiency
Eff=CPM/DPM
Can use efficiency and count rate to determine amount of material present
Amount of light produced per energy deposited in crystal

## Slide 4: Efficiency
Relative efficiency- amount of response from the detector from radiation that has entered the detector volume
Absolute Efficiency- amount of response of the detector from the radiation that is emitted from the source

## Slide 5: Efficiency
What the instrument reads divided by the how much radiation is being emitted
Will be affected by
  Type of radiation
  Energy of radiation
  Geometry of source-detector
  Shape of detector

## Slide 6: Type
Alpha- any detector that has a barrier between the sample and active volume alphas may not have enough energy to penetrate
Beta- see above, energy of beta will determine the probability to penetrate barrier, will change eff from low to high as energy increases

## Slide 7: Types
Gamma-
  thin barrier no problem for gammas, lower energies may be better to interact with detector better than high energy gammas
  Al/metal/plastic protective coverings on some scintillators will block low energy gammas and their eff will be very low
  High energy gammas have low probability of interacting so eff will be low

## Slide 8: Energy
Low energies- hard time penetrating into detector
High energies- easier to penetrate into detector
  but have lower probability to interact for gammas
  Once charged particle get into detector they have high intrinstic (relative) eff

## Slide 9: Shape
Right circular cylinder-
  most popular shape
    Easy to manufacture
    Easy to couple to PMT
Well counters-
  Increased efficiency
Flat thin
  Used to changed particle detection

## Slide 10: Geometry
How is the sample oriented to the detector
  Straight in front
  Off to side
  How far away
What is the make up of the sample
  Solid
    Density
    Layers
  Liquid
    Layers
    Settling

## Slide 11: Sample geometry
Density of sample will determine eff of detector
  Lower density higher eff
  Higher density lower eff
Location
  Directly in front of detectors is best
  Off to the side decreases eff (not bad if you know how much)
  If sample is farther from detector will decrease eff

## Slide 12: Sample make up
Soil
Water
Homogeneous
Heterogeneous

## Slide 13: Standard vs Sample
Standard you use to calibrate detector must be as close to the radionuclide that you are searching for
Standard must simulate the sample
  Density, etc
Static vs Scan efficiencies
  Static -detector is in single point for entire count time
  Scan -detector is moving across the surface

## Slide 14: Release limits
Most release limits are stated in activity per area
  dpm/ 100 cm 2
Need to find out if your instrument can detect down to this level
If your instrument can not detect to at least 10% below the release limit you can not use it for that purpose

## Slide 15: Counting Statistics
MDA
LLD
SD
Count times

## Slide 16: MDA
Minimum Detectable Activity- smallest activity that can be quantified
MDA= (2.71) + 4.64(Rb/T)1/2
---------------------------------------------
E (A/100 cm2)
T= time of background and sample
Rb=background count rate
E= efficiency of detector
A=effective area of detector

## Slide 17: MDA
Background rate =50 cpm
Time =1 minute
Eff= 15 %
Area of probe= 25 square cm
MDA= ?

## Slide 18: LLD
Lower Limit of Detection- smallest activity that can be detected
LLD= 1.64 (Rb/Tb)1/2
-------------------------------------------------------------------------------------------
E (A/100cm2)
Rb = background count rate
Tb = background time
E= efficiency of detector
A= effective area of detector

## Slide 19: LLD
Background rate = 50 cpm
Time =1 minute
Eff= 15 %
Area of probe= 25 square cm
LLD= ?

## Slide 20: Standard Deviation
Standard Deviation- measures the average difference from all number from the average of those numbers
SD= √(Σ(ni-nave)2/(N-1))
ni = number in list
nave= average of number
N = number of numbers in list

## Slide 21: Standard Deviation
Standard Deviation measures the dispersion of information. If the SD is low, the data points are close to the mean. If the SD is high, the data is spread out.
Normal distribution is when
there is an even distribution
around the mean.
Like here:
Notes: The standard deviation is used to designate the variation of the set of numbers. The SD will let the user know if the numbers are tightly grouped around the average or if there is a wider distribution of numbers

## Slide 22: Standard Deviation
With a normal distribution (gaussian distribution)
When the SD is low the variation is low and the curve is tall and peaky (all the data is close to the average)
When the SD is high the variation is high and the curve is low and flat (all the data is spread out and can be very different than the average)
Notes: the standard deviation σ (sigma) is the square root of the average value of (X – μ)2. that is the average value of the differences between all the numbers and the average number of the set.
The Gaussian distribution is named for Carl Friedrich Gauss, who used it to analyze astronomical data, and defined the formula for its probability density function.
For more information please read en.wikipedia.org/wiki/standard_deviation

## Slide 23: Standard Deviation
Number of SDs
  One SD on either side of the mean encompasses 34% of the graph
  One SD on both sides of the graph encompasses 68% of the graph
  2 SD on both sides of the graph encompasses 95 % of the groph
  3 SD on both sides of the graph encompasses 99.8% of the graph
Notes: The number of SDs is used to designate the precision of a measurement. If one can say that their measurments are good to within 2 SDs they are saying that 95% of the time their reading will be correct.

## Slide 24: Detector Confidence
The number of SDs is used to designate the precision of a measurement
If one can say that their measurements are good to within 2 SDs they are saying that 95% of the time their reading will be correct
The higher the SD used the more precise your numbers will be

## Slide 25: SD
SD of net count rate
=√ Rt/Tt+ Rb/Tb
SD of background count
= √ Rb/Tb

## Slide 26: SD
Background = 50 cpm
Background time= 10 minutes
Sample count rate = 500 cpm
Sample time = 5 minutes
SD =

## Slide 27: Count times
Maximization of count times
Ts/Tt-s = √Rs/Rb
Or
Ts/Tb = √Rs/Rb

## Slide 28: Max count times
Sample count rate = 800 cpm
Background count rate= 50 cpm
Sample count time= 5 min
Background count time= 10 min
If you had an hour to count both BG and sample, how long would you count the sample for?

## Slide 29: Questions?