# LAB: An Electrocardiogram Circuit

## Data Acquisition and Measurement

### Acknowledgements

Subject Matter Expert:  Jesus Casas, Austin Community College, Texas

### Purpose

The purpose of this lab is to introduce the student to biological electrical signals and how they can be captured.

### Systems Rationale

A basic understanding of biological signals is important in the troubleshooting of biomedical equipment.  In the case of an Electrocardiogram (ECG) machine, improperly placed electrodes may cause the equipment to appear faulty.  An ECG machine is a complete system with many subsystems.  This lab allows the student to verify each subsystem, as well as the entire system.

### System Concepts

This system covers the following system concepts (signified by an X):

_X_ 	S1. A system can be defined in terms of its functional blocks i.e., a “structured functional unit.”

_X_ 	S2. A system has a purpose, transforms inputs into outputs to achieve a goal.

_X_ 	S3. A system is defined by the flow of materials, energy and information, between its functional units.

_X_ 	S4. A system may be open or closed. In an open system additional inputs are accepted from the environment.

_X_ 	S5. A system is more than the sum of its parts. Individual components can never constitute a system.

_X_ 	S6. A system provides feedback to the operator and services to the user. Some system functions may involve operator action.

_X_ 	S7. Systems have unique problems.

### Learning Outcomes

These LOs maybe found at the SLO tab of the course.  http://www.esyst.org/Courses/Acquisition_Measurement/_delivery/index.php

1.	Draw a block diagram of a modern digital oscilloscope and explain how it works.

2.	Make measurements on analog and digital signals using a digital oscilloscope.

5.	Explain the need for automated test equipment (ATE) systems and state an example case.

7.	Explain the need for and operation of GPIB and LXI systems in ATE systems.

### Prerequisite Knowledge & Skills

The student should have a basic understanding of operational amplifiers and LabView.

### Learning Objectives

The student will be able to state how the electrical activity of the heart occurs.

The student will be able to state some of the safety features found on an ECG machine.

The student will be able to demonstrate the effect a low-pass filter has in an ECG circuit.

The student will be able to state the purpose of making noise the common-mode signal.

### Process Overview

The student will first be introduced to the biological electrical signals of the human heart, then to the circuitry required to capture such signals.  Finally the student will interface the hardware to software (LabView).

### Instructional Setting

This lab has been designed to take place in an Electronics Laboratory with PCs running LabView.

### Time Needed

Instructor Setup/Cleanup:

This will vary depending on the particular lab room arrangement and storage systems available at your institution.

Student Lab Performance:

It should take students approximately 3 hours to work through the entire lab.

### Materials, Equipment, & Supplies Needed

| Item | Quantity |
| --- | --- |
| INA126 (Instrumentation Amplifier) | 1 |
| OP77 or LM741 (Operational Amplifier) | 3 |
| 100 KOhm Resistor | 6 |
| 10 uF Electrolytic Capacitor | 1 |
| .1 uF Capacitor | 1 |
| 5 KOhm Potentiometer | 1 |
| 220 Ohm Resistor | 1 |
| 68 KOhm Resistor | 1 |
| 10 KOhm Resistor | 1 |
| 7805 Voltage Regulator | 1 |
| 9 Volt Battery | 1 |
| 1 MegOhm Resistor | 1 |
| Disposable Electrodes (these can be found on-line) | 3 |
| Electrode leads (these can be found on-line) | 3 |
| ADC0831 | 1 |
| DB25 Male Connector | 1 |
| Oscilloscope | 1 |
| Computer with Parallel Port running LabView 7.0 or greater. | 1 |

### Special Safety Requirements

The ECG circuit must be battery powered only.  Do not use a power supply that is connected to an outlet.

### Assessment & Grading

It is the instructor’s decision as to how students’ learning will be assessed and how this activity is graded.  No specific assessment or grading information is included in this guide.

### Introduction

The human body is very complex with a concert of activity happening at all times.  A very important facet of this activity is the electrical signals that are used to control, regulate, and provide feedback between cells, muscles, organs, and the human brain.  A physician trained in interpreting the electrical signals of a specific area can diagnose and recommend treatment based on the interpretation of those signals.

The human heart, which beats an average of 72 beats per minute, is controlled by electrical signals as well.  By monitoring the electrical activity of a patient’s heart, a heart specialist can determine its condition.

In 1901, Dr. Willem Einthoven created a simple machine that could observe and record the electrical impulses of the heart.  He called the machine an electrokardiogram, or EKG.  The English spelling of electrokardiogram is electrocardiogram, or ECG.  These two names and acronyms are commonly used synonymously.

Typically though, a recording of the heart’s electrical activity from the body surface is called an electrocardiogram, while the instrument itself is referred to as an ECG or EKG machine.  An ECG machine is capable of isolating and amplifying the minute electrical signals of the heart which are typically in the microvolt range

### Introduction to Bio-potentials

Biological cells maintain an electrical potential gradient across the cell membrane due to various ionic pumps on the cell surface.  The potentials are usually less than 100 mV with the inside of the cell being negative with respect to surface of the cell.  Cells are normally considered to be polarized when at rest but can become depolarized when stimulated, temporarily losing the potential difference between the inside of the cell and its outer membrane.  Muscles cells exhibit this depolarization during muscle activity.  This depolarization occurs when gates on the cell membrane allow the passage of ions such as sodium, potassium, and calcium from one side of the membrane to the other.  When this happens, the potential difference between the two sides of the cell membrane diminishes and sometimes even reverses.  When a group or a series of cells depolarizes at the same time or sequentially, an electrical waveform is produced.

The human heart is a large muscle that pumps blood throughout the body.  The heart is composed of four chambers: two atria that receive venous blood and two ventricles that pump arterial blood.

The right side of the heart pumps oxygen depleted blood to the lungs.  After the blood has been oxygenated, it returns to the left side of the heart where it is pumped throughout the body.  The potential generated by the depolarization and re-polarization of the cardiac muscle can be recorded through the proper placement of electrodes on the skin surface.  The depolarization of the muscle cells triggers the contraction.

The heart has a group of cells called the sinoatrial or SA node that are found in the upper wall of the right atrium.  This group of cells is the “pacemaker” for the heart and is responsible for triggering a “heart beat”.  Depolarization of the SA node generates a current that leads to the depolarization of all the other cardiac muscle cells.  The depolarization wave travels from the right atrium to the left atrium causing their contraction.  The wave travels quickly enough that both contract at about the same time.  In the right atria there is another group of cells called the atrioventricular or AV node that will conduct the depolarization down to the ventricles.  A short moment after the atria contract, the ventricles are finally depolarized and their contraction begins.  As soon a contraction is over, repolarization occurs and the muscle awaits another depolarization upon which it will contract once again pushing blood through or out of the heart.

A simplified view of heart activity for one cycle is that SA node causes the atria to contract, forcing blood into the ventricles.  Shortly after the atria contract, the ventricles contract due to the depolarization received from the AV node.  The blood is pushed out of the ventricles and into the aorta and pulmonary artery.  The heart muscles then repolarize and the heart cycle begins again.

The depolarization wave which travels throughout the heart causes a varying potential across the heart itself as it undergoes depolarization, contraction, and repolarization.  The varying potential across the heart is the summation of all the action potentials from the heart, as detected from the surface of the body.

### The ECG Signal

The ECG signal is characterized by six peaks and valleys as seen in the following diagram.  The P wave is due to the depolarization (contracting) of the atria and priming the ventricle with blood.  The QRS is due to depolarization (contracting) of the ventricles, whereby blood is pushed through the pulmonary artery to the lungs and through the aorta to the body.  And the T wave is due to the re-polarization of the ventricles, after which they await another heartbeat.

### Block Diagram for an ECG Circuit

The front-end of any ECG machine is a bio-amplifier.  The term bio-amplifier simply means that it amplifies electrical signals from a biological source.  Typical characteristics are high gain, high-input impedance, high common-mode rejection, low drift, and electrical isolation.

The bio-amplifier for the ECG circuit described in this document consists on an Instrumentation Amplifier, a low-pass filter, and a second stage of amplification.

Because the heart electrical signal is embedded in a “sea” of other stronger electrical signals, an Instrumentation Amplifier is used as the first stage of amplification.  The Instrumentation Amplifier provides high differential gain while providing a high degree of rejection to common-mode signals.

The common-mode signal is usually 60 Hz AC interference and or DC electrode offset potential.  Other noise may be produced from physical movement that changes the skin-electrode interface, muscle activity that introduces electromyographic signals, and electromagnetic interference.  Most of the noise can be eliminated by using an Instrumentation Amplifier with a high Common-Mode Rejection Ration (CMRR).

Even though most noise signals may be eliminated by the Instrumentation Amplifier, there may still be some unwanted signals present at the output of the Instrumentation Amplifier.  It is most likely that these unwanted signals are at higher frequencies than that of the heart’s beat.  Therefore a low-pass filter is used to eliminated the high frequency noise signals and allow the low frequency heart signal to pass through for more amplification.

Taking the average heart rate of 72 beats per minute when at rest and allowing for a heart rate of three times as much when the heart is exercised, a cut-off frequency for the low-pass filter can be arrived at.

The output of the bio-amplifier will then be fed to an Analog-to-Digital Converter, whose output will then be processed by a microprocessor for viewing of the signal.  The software running on the microprocessor will be LabView.

### Safety

Connecting an electrical circuit across a patient’s heart always carries some inherent danger with it.  This is especially true when the circuit is powered from an AC outlet.  To provide a degree of safety, the following have been used in the ECG circuit shown in this document:

Diode Clamping – The Instrumentation Amplifier has clamping diodes at both of its inputs.  This prohibits the voltage at the inputs from exceeding +.6 Volts or -.6 Volts (in respect to the Instrumentation Amplifier ground).

Battery Operation – The entire ECG circuit is powered by a single 9 Volt battery.  This provides a DC source and eliminates an outlet-driven power supply which could potentially allow 110VAC to come directly across the patient’s heart in the event of a catastrophic failure.

Input Resistors – Even though the Instrumentation Amplifier has an input impedance of approximately 60 Meg Ohms, input resistors are used to safely limit the current that could potentially be made to flow out of the inputs in the event of a catastrophic failure.

Further safety could be incorporated by using an isolation amplifier or opto-isolators.

### Electrode Placement

The bio-amplifier takes inputs from self-sticking electrodes that are attached to the patient’s body.  Proper placement of the electrodes is crucial in obtaining a good electrical signal.  The ECG circuit shown in this document requires that three electrodes be used.  The following diagrams indicate the locations that can be used.

By placing the electrodes as shown, the electrical activity of the heart can be captured.  The following diagram illustrates the vector associated with the depolarization and repolarization of the heart muscles.

### Microprocessor (Computer) Interface

The output of the bio-amplifier will be converted into a digital signal by an 8-bit Successive Approximation Analog-to-Digital Converter (ADC).  The ADC output will be serial in nature and will be connected to the parallel port of a computer running LabView.  LabView will allow the viewing of the ECG signal in a graphical format.

The following parallel-port interface circuit employs an ADC0831.  This circuit was chosen because of its low I/O requirement.

### The Bio-Amplifier Circuit

The first stage of the bio-amplifier is the Instrumentation Amplifier which uses a gain resistor of 220 Ohms.  Even though a lesser resistance would have provided greater gain, it would have also introduced far more noise that would out weigh any benefit of additional gain to the wanted signal.

The INA126 requires a bi-polar power supply.  Since the circuit is operated from a single 9 Volt battery, an artificial ground was created by OA-1 (Operational Amplifier).  OA-1 provides an output voltage of +4.5 Volts.  The INA126 uses this voltage as its “ground”.  From the perspective of the INA126, it is being fed + 4.5 Volts for its bi-polar power supply.  From the perspective of circuit ground, the output of the INA126 is biased at +4.5 Volts.

The output of the INA126 is then routed through an adjustable low-pass filter comprised of a 5 KOhm potentiometer and a 10uF capacitor.

The AC portion of the filtered signal is then passed through for further amplification by OA-3.  OA-3 is configured as a Non-Inverting Amplifier.  Since the output of OA-3 is to be fed to the ADC and the ADC is configured for conversion between 0 and +5 Volts, the output of OA-3 is therefore biased at +2.5 Volts.  This +2.5 Volt bias voltage is provided by OA-2.

OA-2 provides an output voltage of +2.5 Volts.  OA-3 uses this voltage as its “ground”.  From the perspective of OA-3, it is being fed +2.5 Volts for its bi-polar power supply.  From the perspective of circuit ground, the output of OA-3 is biased at +2.5 Volts.

### Lab Procedure

1.	Calculate the gain of the Instrumentation Amplifier in the Bio-Amplifier circuit.

The gain is __________.

2.	In the event of a catastrophic failure and if 110V were seen across the inputs of the Instrumentation Amplifier, calculate the maximum current that could flow through the electrodes?

The maximum current that could flow through the electrodes is ______________.

3.	Calculate the gain of the Non-Inverting amplifier comprised of OA-3.

The gain is __________.

4.	Assemble the Bio-Amplifier circuit.

5.	Supply +9 Volts to the Bio-Amplifier circuit.

Warning:  Use only a 9 Volt battery.  Do not use a power source that is connected to an outlet.

6.	Verify the output of OA-1 is +4.5 Volts.

Voltage verified:  _____ (check-off)

7.	Verify the output of OA-2 is +2.5 Volts.

Voltage verified:  _____ (check-off)

8.	Verify the output of OA-3 is biased at +2.5 Volts.

Voltage verified:  _____ (check-off)

9.	Connect the input of an Oscilloscope to the output of the INA126.  Set the amplitude of the Oscilloscope to .5 Volts per division and the time base to 500 mSec per division.

10.	Apply electrode gel to three electrodes.

11.	Using one of the electrode-placement diagrams, place the electrodes on a “patient”.

12.	Connect the “patient” to the Bio-Amplifier through use of electrode leads.

13.	Monitor the Oscilloscope, a heart-beat signal should be seen.  Please note that some noise is expected to be seen as well.

14.	Connect the Oscilloscope to the output of OA-3, a heart-beat signal should be seen.  Adjust the 5KOhm potentiometer of the low-pass filter so as to adjust the cut-off frequency of the filter.  Adjust the potentiometer so that the signal is relatively noise free and still resembles the waveform of the ECG signal diagram shown earlier.

15.	With the Oscilloscope still connected to the output of OA-3, measure the Vpp of the heart-beat signal.

Vpp is ______________ Volts.

16.	Having calculated the gain of the Non-Inverting amplifier (Step 3) and measured the Vpp of the heart-beat signal (Step 15), calculate the Vpp of the signal at the input of the Non-Inverting amplifier.

Vpp is ______________ Volts.

17.	Connect the Oscilloscope to the output of the INA126.  Measure the Vpp of the heart-beat signal.

Vpp is ______________ Volts.

18.	Calculate the voltage gain (Av) of the low-pass filter using the results of Step 16 and 17.

Av for the low-pass filter is __________.

19.	Having calculated the gain of the INA126 (Step 1) and measured the Vpp of the heart-beat signal at its output (Step 17), calculate the Vpp of the electrical signal the electrodes are sampling.

Vpp is ______________ uVolts.

20.	Calculate the overall voltage gain for the Bio-Amplifier.

Av for the Bio-Amplifier is __________.

21.	Disconnect the “patient” from the Bio-Amplifier.

22.	Assemble the A/D circuit making the proper connections to the Bio-Amplifier and to the Parallel Port of a PC running LabView.  The A/D circuit will get its power from the +5 Volt regulator used in the Bio-Amplifier.

23.	Run the LabView program titled “ECG_ADC0831.vi”

24.	Connect the “patient” to the Bio-Amplifier once again.  A heart-beat signal should be seen.  The following is a picture of the VI’s front panel displaying a heart-beat signal.

25.	The program has a time delay (in mSec) between samples as indicated in the following picture.

Adjust this value to 1 mSec and run the LabView program again.  Since the program is sampling at a faster rate, what is the result of the signal seen?  Is the ECG signal still present?  Is there more “noise” seen?

________________________________________________________________________________________________________________________________

26.	Adjust the time delay value to 50 mSec.  Since the program is now sampling at a slower rate, what is the result of the signal seen?  Is the waveform of the ECG signal still like the one seen in Step 24?

________________________________________________________________________________________________________________________________

End of Lab