Student # __________      Group # __________                            TA _______________

Vacuum Equipment: Components and Systems Part 2A

Objective: It is time to put it all together.  From the available material provided with the MKS trainer, design and build a vacuum system that is capable of reaching pressures of  5 x 10-5 Torr.  This means the system has to reach and maintain the required pressure in a reasonable time.  The system must be rugged and reliable with little down time.  To accomplish this, proper construction, operating procedures, and maintenance need to be considered.

Background: Through lectures and labs you have learned about the importance of vacuum.  It creates a clean environment, lowers the molecular density of gasses, provides the necessary condition to create plasma, can be used to control the rate of reactions.  By understanding the basic concepts and components of a vacuum system and its basic software, the student will be able to operate equipment with confidence and be able to better understand and design an operational system.  When things do go wrong in a vacuum system, the student will be able to identify and correct the problem, whether it be equipment or product-related.

Experiment:  As a group, perform the following task:

Design a simple operational vacuum system that is capable of reaching a pressure of 5 x 10-5 Torr.

Assemble the designed vacuum system.

Write an operation procedure (including start up and shut down).

Determine if the system is tight (leak free).  Find and repair any leaks present.

Procedure:

Design and draw a block diagram of a vacuum system.

The system must be able to achieve 5 x 10-5 torr.

The diagram must include the appropriate pumps, vents, and valves.

All items must be labeled and include a description.

Write a step-by-step procedure for operating the system and discuss with the instructor.

- Note: Do not proceed without your instructor’s approval.

Bring the system to base rough pressure.

Bring the system to base hi-vac pressure.

Write a step-by-step procedure for venting the chamber and shutting down the system.

- Note: Do not proceed without your instructor’s approval.

Vent chamber and shut down system.

- Block Diagram

- Include valves and gauges.

- All items must be labeled and described.

PUMP DOWN PROCEDURE     SHUT DOWN PROCEDURE

Vacuum Equipment: Components and Systems Part 2B

Objective:  The objectives of this lab are to help the student better understand how vacuum systems operate.  Identify a leak, determine the type of leak (real or virtual), isolate, locate and repair the leak.

Background:  As you have discovered, the MKS is a real working vacuum system.  From the previous exercise, it was learned that the up leak test can be used to assess the total amount of leakage in a vacuum system.  However, this method will not serve to actually locate the leak.  In order to locate a leak, a common method is to spray each likely area with a probe gas, such as helium.  To determine the location of the leak, a detector that is sensitive to the probe gas must be placed downstream of the suspected location.  An indication (audio or visual) will be given by the detector when the probe gas is transported to it.  For detecting very fine leaks, the standard method is to use a helium spectrometer leak detector.  In our system, we will use the gas sensitivity of the residual gas analyzer (RGA) to perform the same function.

Figure I: Schematic of a quadrupole analyzer1

In order for the RGA to be utilized as a leak detector, the system must be able to reach a pressure of at least 5 x 10-4 Torr.  The RGA utilizes an open ion source (OIS) quadrupole analyzer to detect helium in the vacuum chamber.  The ionizing filament and collecting anode are “open” to the inside of the vacuum system (Figure I).  Positive ions emerge from the filament and are focused down the four rod tuning system.  When helium enters the vacuum system, through a leak, it is ionized at the filament.  Radio frequencies are set to form a stable helium ion path down the rod system to the controller.  Gas ions with masses smaller or larger than the mass of helium will not form a steady ion stream and are collected by the quadrupole rods.  Helium is then detected by the collector.

Experiment: Once the system has reached its base rough pressure, assess the presence of a real or virtual leak by performing an up leak test.  After performing the test and determining if there is a leak, return the system to its base pressure.

Place the helium applicator nozzle gently against the suspected leak and begin spraying with helium.

Observe the RGA computer interface.  Be slow, and deliberate.  Be patient!  Give the gas time to reach the analyzer.

If a leak is not located, proceed to the next area.  In general, you should start farthest from the RGA.  Also, remember the nature of the gas.  Helium is lighter than air, so do not begin at the bottom.  Work from the top down.

When all leaks have been discovered.  Vent the system, repair the leaks, and reassemble.

Bring the unit back to base rough pressure and retest.

References: Harris, N. Modern Vacuum Practice 3rd Ed.. McGraw-Hill 2007

Questions:

1. Leak test before repair: Create an Excel chart for this test.  Be sure to include the following:

Base Pressure: _______    Final Pressure: ___________

Sample Period: __________   Leak Rate: ____________

2. Leak test after repair: Create an Excel chart for this test.  Be sure to include the following:

Base Pressure: _______    Final Pressure: ___________

Sample Period: __________   Leak Rate: ____________

3. Why is helium used for leak testing?

4. What type of leak was fixed in this exercise and explain (real or virtual)?

5.  Research and summarize one reference pertaining to helium leak detection.

6.  Research and summarize an alternative to helium leak detection.

Updated: 6/24/10

PAGE

© 2010 The Pennsylvania State University