ACADs (08-006)  Covered

| 3.2.5.10 | 3.3.14.2 | 3.3.14.3 | 3.3.14.5 | 3.3.14.7.2 | 3.3.14.7.3 | 3.3.14.7.8 |
| --- | --- | --- | --- | --- | --- | --- |
| 3.3.14.7.9 | 3.3.14.7.10 |  |  |  |  |  |

Keywords

Byproduct material, source material, controller, custodian, ORM, sealed source, exempt source, special nuclear material license, controller responsibilities, custodian responsibilities,  receipt, regulation, distribution, inventory leak test requirements, disposal authorization.

Description

This document provides references to materials, instructor notes and a link to an accompanying Powerpoint presentation designed to teach a lesson on radioactive materials handling and accountability.

Supporting Material

Radioactive Materials Handling and Accountability

NUCLEAR TRAINING

TRAINING MATERIALS COVERSHEET

Please TAB to gray areas

| RADIOLOGICAL PROTECTION TECHNICIAN INITIAL TRAINING |  |
| --- | --- |
| PROGRAM FUNDAMENTALS TRAINING HPT001 |  |
| COURSE rADIOACTIVE MATERIALS HANDLING AND ACCOUNTABILITY | COURSE NO. HPT001.310 |
| LESSON TITLE | LESSON PLAN NO. |

| INPO ACCREDITED | YES | X |  | NO |  |
| --- | --- | --- | --- | --- | --- |
| MULTIPLE SITES AFFECTED | YES | X |  | NO |  |

| PREPARED BY Ralph G. Wallace/Brian K. Fike | - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Signature / Date |
| --- | --- |
| PROCESS REVIEW David L. Stewart | - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Signature / Date |
| LEAD INSTRUCTOR/PROGRAM MGR. REVIEW R. L. Coleman | - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Signature / Date |
| PLANT CONCURRENCE | - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Signature / Date |
| TVAN CONCURRENCE (If applicable) | - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - - Signature / Date |
|  | BFN SQN WBN CORP |

| Receipt Inspection and Distribution: |  |  |
| --- | --- | --- |
|  | Training Materials Coordinator | / Date |

| Standardized Training Material |  |
| --- | --- |
| Copies to: |  |

###### TVA 40385 [NP 6-2003] Page 1 of 2

| NUCLEAR TRAINING REVISION/USAGE LOG |  |  |  |  |
| --- | --- | --- | --- | --- |
| Rev. # | Description of Changes | Date | Pages Affected | Reviewed By |
| 0 | Initial Issue | TBD | ALL | Ralph G. Wallace/ Brian K. Fike |

###### TVA 40385 [NP 6-2003] Page 2 of 2

I.	PROGRAM:	Radiological Protection Technician Initial Training

II.	COURSE:	Fundamentals Training

III.	TITLE:	Radioactive Material Handling and Accountability

IV.	LENGTH OF LESSON:	2-3 hours

V.	TRAINING OBJECTIVES

A.	Terminal Objective

Upon completion of this module, participants will demonstrate knowledge of the handling and accountability of radioactive material as it pertains to Radiological Control.  A score of > 80% must be achieved on a written examination.

B.	Enabling Objectives			

Standards and conditions apply to all enabling objectives.  They include the training participant’s ability to utilize, under the examination ground rules (i.e. without the use of training materials or outside assistance), the information presented in this lesson plan.

1.	In the context of Radioactive Materials (RAM), define the following:

a.	Byproduct and Source Material (BSM) Controller

b.	Byproduct and Source Material (BSM) Custodian

c.	Byproduct Material

d.	Item Control Area (ICA)

e.	NORM

f.	Physical Inventory

g.	Sealed Source

h.	Source

i.	Exempt Source

j.	Non-Exempt Source

k.	Source Material

l.	Special Nuclear Material (SNM)

m.	User

2.	Identify the document that contains the plant operating license, which also covers the possession and use of BSM and SNM at the site.

3.	Identify the TVAN procedure that addresses the control of BSM.

4.	State the responsibilities of the BSM Controller.

5.	State the responsibilities of the BSM Custodian.

6.	Tell the responsibilities of the SNM Custodian.

7.	Identify the organization that is responsible for receiving materials, including radioactive materials, at each site.

8.	Identify the time period within which packages of radioactive materials received with a Radioactive White I, Yellow II, or Yellow III label must be surveyed.

9.	Identify the NRC regulation that defines exempt concentrations and exempt quantities.

10.	Identify the NRC regulation covering the licensing of persons who allowed to manufacture and distribute exempt quantity sealed sources.

11.	State the requirements for conducting an inventory of BSM.

12.	Tell the actions taken in the event of a lost, stolen, or missing licensed sealed source.

13.	Identify the requirements for conducting leak tests of licensed sealed BSM sources.

14.	Sate the general procedures used in performing leak tests of licensed sealed BSM sources.

15.	Calculate the activity on a leak test smear and identify the activity level on a smear that would indicate that a sealed source is leaking.

16.	Identify the site organization responsible for the disposal of unwanted BSM.

VI.	TRAINING AIDS

A.	Whiteboard with markers.

B.	Projector and Screen

C.	Power Point Presentation of Key Points

D.	Laser Pointer (optional)

- TRAINING MATERIALS:

A.	Appendices

1.	Handouts

a.	HO-1 – Enabling Objectives

b.	HO-2 – Terms and Definitions

c.	HO-3 – Responsibilities

d.	HO-4 – Radiological Limits for the Receipt of Radioactive Material

e.	HO-5 – Nonexempt Source User's Responsibilities – Instruction Sheet

f.	HO-6 – Exempt BSM Sources and Devices

g.	HO-7 – Attachment 4

h.	HO-8 – Attachment 5

B.	Attachments

1.	Power Point slide show Radioactive Material Handling and Accountability

a.	TP-01 – Radioactive Material Handling and Accountability

b.	TP-02 – Enabling Objectives

c.	TP-03 – Enabling Objectives

d.	TP-04 – Enabling Objectives

e.	TP-05 – Enabling Objectives

f.	TP-06 – Enabling Objectives

g.	TP-07 – Basis Documents

h.	TP-08 – BSM Controller

i.	TP-09 – BSM Custodian

j.	TP-10 – SNM Custodian

k.	TP-11 – Procuring Radioactive Materials

l.	TP-12 – Receiving Radioactive Materials

m.	TP-13 – Inspect Packages for Damage

n.	TP-14 – Receipt of Special Nuclear Materials

o.	TP-15 – BSM Storage Accountability & Use

p.	TP-16 – Assignment of Sources

q.	TP-17 – Exempt Sources

r.	TP-18 – Control of BSM

s.	TP-19 – BSM Source Inventory

t.	TP-20 – “Found” BSM Sources

u.	TP-21 – “Lost” BSM Sources

v.	TP-22 – Leak Testing BSM

w.	TP-23 – General Guidelines for Conducting Leak Tests

x.	TP-24 – Evaluating Leak Test Smears

y.	TP-25 – SNM Inventory

z.	TP-26 – Transfer and Disposal of RAM

aa.	TP-27 – QA Records

bb.	TP-28 – Summary

2.	OE17854 – Wolf Creek – “Unanticipated Fixed Contamination Discovered on Vendor Source Packaging.”  At web site: File://C:\WINDOWS\Temp\InpoReader915870.htm

3.	OE16021 – River Bend Station – “Source Misplaced Subsequent to Fuel Sipping Equipment Check.”  At web site:  File://C:\WINDOWS\Temp\Inpo115984.htm

4.	10 CFR 30.70, Schedule A - Exempt Concentrations.  At web site:  http://www.nrc.gov/reading-rm/doc-collections/cfr/part030/part030-0070.html

5.	10 CFR 30.71, Schedule B – Exempt Quantities.  At web site:  http://www.nrc.gov/reading-rm/doc-collections/cfr/part030/part030-0071.html

6.	Licensee Event Report 254-94009, 9/15/1994, “Failure to Maintain a Current Inventory of Radioactive Sources.  At web site:  http//www.inpo.org/databases/events/94/254_940823_1.htm

VIII.	REFERENCES:

A.	ACAD 93-008, “Guidelines for Training and Qualification of Radiological Protection Technicians,” National Academy For Nuclear Training, August 1993.

B.	Code of Federal Regulations Title 10, Part 20, “Standards for Protection Against Radiation,” U. S. Government Printing Office, Washington, 2003.

C.	Code of Federal Regulations Title 10, Part 30, “Rules of General Applicability to Domestic Licensing of Byproduct Material,” U. S. Government Printing Office, Washington, 2003.

D.	Code of Federal Regulations Title 10, Part 34, “Licenses for Industrial Radiography and Radiation Safety Requirements for Industrial Radiography Operations,” U. S. Government Printing Office, Washington, 2003.

E.	Code of Federal Regulations Title 10, Part 40, “Domestic Licensing of Source Material,” U. S. Government Printing Office, Washington, 2003.

F.	Code of Federal Regulations Title 10, Part 70, “Domestic Licensing of Special Nuclear Material,” U. S. Government Printing Office, Washington, 2004.

G.	Code of Federal Regulations Title 10, Part 71, “Packaging and Transportation of Radioactive Material,” U. S. Government Printing Office, Washington, 2004.

H.	Code of Federal Regulations Title 10, Part 74, “General Control and Accounting of Special Nuclear Material,” U. S. Government Printing Office, Washington, 2004.

I.	Code of Federal Regulations Title 49, Part 173, “Shippers-General Requirements for Shipments and Packagings,” U. S. Government Printing Office, Washington, 2003

J.	TVAN Standard Programs and Processes SPP-5.6, “Controlling Byproduct and Source Material,” Revision 4, 8/18/2003.

K.	TVAN Standard Programs and Processes SPP-5.8, “Special Nuclear Material Control,” Revision 5, 3/19/2004.

L.	TVAN Standard Programs and Processes SPP-5.9, “Radiological Control and Radioactive Material Shipment Augmented Quality Assurance Program, 
Revision 5, 10/8/2003.

M.	BFN RCI-7, “Receipt of Radioactive Materials,” Revision 16, 9/12/2003.

N.	SQN RCI-06, “Receipt of Radioactive Materials,” Revision 17, 3/14/2003.

O.	SQN RCI-17, “Control of Byproduct and Source Material,” Revision 15, 9/6/2002.

P.	SQN RCI-21, “Control of Radioactive Material,” Revision 11, 2/1/2002.

Q.	SQN 0-SI-RCI-000-056.0, “Byproduct Material Inventory and Sealed Source Leak Test,” Revision 7, 11/24/2003.

R.	WBN RCI-103, “Radioactive Material Control,” Revision 22, 1/17/2003.

S.	WBN RCI-127, “Byproduct and Source Material Control,” Revision 4, 2/12/2003.

T.	WBN Technical Requirement Instruction 0-TRI-0-3, “Leak Testing of Sealed Sources,” Revision 3, 12/14/1999.

U.	BFN Technical Specifications, Unit One, Unit Two, and Unit Three.

V.	SQN Technical Specifications, Unit One and Unit Two.

W.	WBN Technical Specifications, Unit One.

X.	http://www.crcpd.org/links/Prescreened_Links/Rad_material_index.asp

Y.	http://npj.goinfo.com/NPJMain.nsf/0/552c47153a8f9b3886256a020077df7d?
OpenDocument.  (This excerpt is from the remarks of Richard A. Meserve, Chairman, U.S. Nuclear Regulatory Commission, at the IAEA Conference on Regulatory Control of Radioactive Sources on December 11, 2000 in Buenos Aires, Argentina.)

IX.	INTRODUCTION:

Radioactive material, both that which occurs naturally in the environment and that which is man-made in a reactor or accelerator, is used in industry and medicine for a myriad of functions, as well as being used in the production of electric power.  It can be used to diagnose and treat diseases in humans and animals, detect flaws in structures, airplanes, and pipelines, gauge the levels of materials in tanks, and sterilize medical supplies, to name a few.

Radioactive material use is regulated at the federal and state levels of government. Although these materials are carefully regulated, occasionally radioactive sources are lost or stolen and may wind up in scrap yards or simply become “orphaned.”  (Reference X)

The number of operating nuclear power reactors around the world is relatively small—approximately 440 reactors, of which 103 are in the United States—but they attract close attention.  In contrast to the 103 licensed nuclear power plants in the U.S., there are about 150,000 licensees for radioactive materials in the U.S. and about 2.0 million devices containing radioactive sources in use by licensees.  If control of radioactive sources is lost, these devices can constitute a significant hazard to the public.  Although the cumulative impacts from an event involving a source may be less than those arising from a reactor accident, the likelihood of an event is larger.  Based on the past experience and the large number of sources, (it is) concluded that the likelihood of serious radiation exposure of a member of the general public is larger from radioactive sources than from civilian reactors.  (Reference Y)

With these thoughts in mind, this lesson plan discusses the controls TVA has placed over the use of radioactive materials and sources at its facilities, especially its nuclear power plants.  It is incumbent on each one of us to do everything we can to ensure that the radioactive materials in our control are used with the utmost care and in accordance with applicable procedures.

X.	TERMS AND DEFINITIONS                                                                                     Obj-1

1.	Byproduct and Source Material (BSM) Controller – That individual within the site RADCON organization or facility RSO organization that performs the BSM source control process.

2.	BSM Custodian – The individual who receives custody of BSM sources for control and use in approved facility activities.

3.	Byproduct Material – Any radioactive material (except SNM) yielded in, or made radioactive by exposure to the radiation incident to, the process of producing or using SNM.

4.	Item Control Area (ICA) – A defined area within a licensed facility with physical boundary delineation wherein Special Nuclear Material (SNM) is kept, subject to material control and accounting practices such that, at any time, an item count and related quantities of SNM can be obtained from the records for the SNM located within the area.

5.	Naturally Occurring Radioactive Material (NORM) – Radioactive material which occurs in nature and is not the product of nuclear fission or a nuclear accelerator.

6.	Physical Inventory – A piece by piece accounting of BSM sources supported by formal documentation.  Acceptable methods of performing the inventory consist of either direct observation of the source or verification by instrument response or other indirect means.  Consolidated sources maintained in sealed containers may be inventoried by ensuring the container is in its assigned location and the seal is unbroken.

7.	Sealed Source – Radioactive material that is encased in a capsule designed to prevent leakage or escape of the material under normal conditions of use and leak tested per facility licenses.

8.	Source, General – Any BSM with known isotopes and activity manufactured for the purpose of measuring, checking, calibrating, or controlling processes quantitatively or qualitatively. These sources include but are not limited to:

A.	Sources in field monitoring equipment to verify operability.

B.	Sources in radiation monitors to maintain a threshold of sensitivity.

C.	Spent uranium slabs used to determine beta response and correction factors for portable monitoring instrumentation.

D.	Sources used to calibrate and response-check field monitoring equipment (portable and fixed).

E.	Liquid standards and liquids and gases used to calibrate and verify calibration of laboratory counting and analyzing equipment.

F.	Sources used for radiographic operations.

9.	Source, Exempt (Non-Licensed) – Those BSM sources exempted from licensing requirements for the user by 10CFR30, 31, and 40. In general, the requirements of 10CFR19, 20, 30, 31, 32, 33, 34, 35, 39, and 40 do not have to be met unless specifically stated in the applicable section. Only sources received from distributors specifically licensed in accordance with 10 CFR 32 or agreement state requirements to distribute license exempt sources, may be considered exempt. Requirements communicated by the sources manufacturer with the “exempt” material are applicable.

10.	Source, Non-Exempt (Licensed) – Any BSM that is not specifically exempted from licensing requirements by the "EXEMPTIONS" section of 10CFR30, parts 30.11 through 30.20, all of 10CFR31, and the "EXEMPTIONS" section of 10CFR40, parts 40.11 through 40.14. A Specific License is required for BSM which does not meet the above exceptions. In general, the requirements of 10CFR19, 20, 30, 31, 32, 33, 34, 35, 39, and 40 must be met.

11.	Source Material – Uranium or thorium, or any combination thereof, in any physical or chemical form, or ores that contain by weight one-twentieth of one percent (0.05%) or more of uranium, thorium, or any combination thereof.  Source material does not include SNM.

12.	Special Nuclear Material (SNM) – Plutonium, uranium 233, uranium enriched in the isotope 233 or in the isotope 235, and any other material which the Nuclear Regulatory Commission, pursuant to the provisions of Section 51 of the Atomic Energy Act, determines to be SNM or any material artificially enriched by any of the foregoing; but does not include source material.

13.	User – An individual who is issued, controls, and uses BSM in performing approved facility activities.

| A. Requirements and Bases | TP-1, 2, 3, 4, 5, & 6 HO-01 & 02 |
| --- | --- |
| 1. A number of NRC regulations address the controls and requirements for handling radioactive materials. These regulations include: | TP-7 |
| a. 10 CFR 20, “Standards for Protection Against Radiation.” |  |
| b. 10 CFR 30, “Rules of General Applicability to Domestic Licensing of Byproduct Material.” |  |
| c. 10 CFR 34, “Licenses for Industrial Radiography and Radiation Safety Requirements for Industrial Radiography Operations.” |  |
| d. 10 CFR 40, “Domestic Licensing of Source Material.” |  |
| e. 10 CFR 70, “Domestic Licensing of Special Nuclear Material.” |  |
| f. 10 CFR 71, “Packaging and Transportation of Radioactive Material.” |  |
| g. 10 CFR 74, “General Control and Accounting of Special Nuclear Material.” |  |
| 2. The authorization from the NRC for the plants to possess byproduct, source, and special nuclear material sources and/or materials is contained in each plant’s operating license, which is a part of the respective plant Technical Specifications. | OBJ.-2 |
| 3. TVA procedures documents that address the control of radioactive materials (RAM) include: |  |
| a. SPP-5.6, “Controlling Byproduct and Source Material.” This procedure provides the Byproduct and Source Material (BSM) source control process for TVAN facilities. It provides for the receipt of radioactive material and the control and accountability of BSM licensed sources. | OBJ.-3 |
| b. SPP-5.8, “Special Nuclear Material Control.” This procedure establishes the administrative controls for management of special nuclear material (SNM). |  |
| c. SPP-5.9, “Radiological Control and Radioactive Material Shipment Augmented Quality Assurance Program.” This procedure describes the Augmented Quality Assurance Program for Radiological Control and Radioactive Material Shipment (RC&RMS). In relation to the control of radioactive materials, it addresses the handling of records generated in the process. |  |
| d. In addition, a number of specific plant instructions (see the reference list) detail procedures for performing specific activities relating to the control of Radioactive Material (RAM). | Error Prevention! Because of the potential for exposure, it is important that procedures be followed completely. |
| B. General |  |
| 1. The BSM Controller is the individual within the site RADCON organization that performs the BSM source control process. He/she has the overall responsibility for and directs the facility BSM Program. | TP-8, OBJ.-4 |
| 2. The BSM Controller may transfer custody of specific RAM sources to other facility groups who have applications for their use. The BSM Custodian in the user group is responsibility for the control, use, and accountability of the sources in his/her custody. | TP-9 OBJ.-5 |

| 3. Each site shall designate a Special Nuclear Materials Custodian (SNMC) and alternates from Reactor Engineering. The SNMC is responsible for the management of all SNM, including: | TP-10 OBJ.-6 |
| --- | --- |
| a. Supply/Receipt Control; |  |
| b. Internal Control; |  |
| c. Isotopic Inventories; |  |
| d. Shipping Control. |  |
| 4. Summary of Responsibilities | HO-03 |
| a. RADCON Manager |  |
| (1) Implement applicable procedures. |  |
| (2) Ensure compliance with Technical Specifications. |  |
| (3) Designate a BSM Controller. |  |
| (4) Provide technicians to assist the Controller. |  |
| b. BSM Controller | OBJ.-4 |
| (1) Maintains the BSM source control program. |  |
| (2) Maintains copies of receipt, transfer, disposal, inventory, and leak test records as applicable in an organized, readily accessible file. |  |
| c. BSM Custodians | OBJ.-5 |
| (1) Controls BSM sources assigned to the Custodian. |  |
| (2) Provides and maintains a storage area to which only the Custodian or designees have access. |  |
| (3) Keeps their inventory and record of issuance to individual users in an organized, readily accessible file. |  |
| d. SNM Custodian | OBJ.-6 |
| (1) Manages and controls all SNM. |  |
| (2) Maintains copies of receipt, transfer, disposal, inventory, and leak test records relating to SNM in an organized, readily accessible file. |  |
| e. User |  |
| (1) Maintains positive control over all BSM sources assigned to the user. |  |
| (2) Uses good RADCON and safety practices when using nonexempt BSM sources. | Includes the application of appropriate Error Prevention Tools. |
| C. Procuring and Receiving Radioactive Materials. | TP-11 |
| 1. Requests for the purchase of BSM may be made by the BSM Custodian or User Groups. |  |
| a. Material requests for BSM are reviewed by the RADCON Superintendent or the BSM Controller to ensure that: |  |
| (1) Materials are actually required for the planned activities, |  |
| (2) Materials are not available in existing inventories, and |  |
| (3) The Custodian/User Group has adequate radiological controls for storage and use of the sources. |  |
| b. The RADCON Superintendent or the BSM Controller approves the BSM material request. |  |
| 2. The SNMC approves all procurement of nonfuel SNM. |  |

| 3. Receiving Radioactive Materials | TP-12 |
| --- | --- |
| a. Picking up, receiving, and opening packages containing BSM shall meet the requirements of 10 CFR 20.1906. Because packages containing BSM may not always be clearly marked for easy identification of type and quantity, methods for the immediate identification of all BSM received shall be established by receiving organizations to ensure these requirements are met. |  |
| b. Arrangements shall be made to receive all packages of BSM when they are offered for delivery, or to take possession of all packages of BSM expeditiously when notification of the package's arrival is received |  |
| c. Packages containing Radioactive materials are normally received by Site Nuclear Stores at the receiving warehouse; however, the BSM Controller may approve alternate receiving locations and/or methods if adequate provisions are made to ensure all receipt requirements and radiological safety precautions are met. An example of this would be receipt of whole body counter sources via the US Mail and delivered directly to the Dosimetry Laboratory. | OBJ.-7 |
| d. The receiver inspects the package containing the BSM prior to removing the package from the transport vehicle. WARNING: IF THE PACKAGE SHOWS OBVIOUS SIGNS OF DAMAGE OR LEAKAGE, A RADIATION OR CONTAMINATION HAZARD MAY EXIST. |  |
| e. If the package shows signs of damage or leakage, the receiver shall notify RADCON immediately. The package shall be left on the transport vehicle and the driver and vehicle shall be detained until RADCON technicians have arrived to assess the situation and establish radiological controls. |  |
| Transparency # 11 is a photo of a packaged received at TVA several years ago. Fortunately, it did not contain radioactive material, but think of the potential consequences if it had. | TP-13 |
| f. If the package is intact and has arrived on a non-sole-use vehicle, the receiver may remove the package from the vehicle and allow the vehicle to leave site. Sole-use vehicles must be surveyed by RADCON before they can be returned to service. |  |
| g. The receiver shall notify RADCON as soon as possible after receiving any package containing radioactive materials. If the package contains special nuclear material, the SNM Custodian shall also be notified. |  |
| h. All packages containing radioactive materials are normally opened in the presence of a RADCON representative. If the item is identified as externally contaminated, it must be opened in a Radiologically Controlled Area (RCA). |  |
| i. All packages containing radioactive material shall be monitored for contamination and radiation levels upon receipt. For packages meeting the conditions outlined below, the survey must be conducted within 3 hours after receipt during normal working hours, or no later than 3 hours after the beginning of the next working day if received after working hours. | OBJ.-8 |
| (1) All packages with a Radioactive White I, Yellow II, or Yellow III label; |  |
| (2) All packages containing more than a Type A quantity and having a DOT Radioactive White I, Yellow II, or Yellow III label; |  |
| (3) Any radioactive material package showing evidence of degradation. |  |
| OE17854 – Wolf Creek – A radioactive source package from a vendor was found to be contaminated with fixed contamination on the inside of the package. The contamination was attributed to less than adequate vendor contamination control techniques when packing the source (human performance, inadequate procedures). Error prevention tools; Procedure adherence, Self-checking, Attention to detail. | Review OE17854, Attachment 2 |
| j. Handout # 04 presents the limits for contamination and radiation levels from 10 CFR 71 and 49 CFR 173. If these limits are exceeded, the following actions are taken: | HO-04 |
| (1) Document the occurrence and immediately notify the RADCON Superintendent or his designee. |  |
| (2) Notify the Shift Manager who will determine if the occurrence is reportable. (If the receipt survey indicates removable contamination or radiation levels in excess of those listed in 10 CFR 20.1906 (d), TVAN must immediately notify the final delivery carrier and the NRC Operations Center by telephone). |  |
| k. Requirements for the receipt of special nuclear material. | TP-14 |
| (1) If the material is shipped in a container without a tamper indicating seal, the SNM Custodian must open the container within two working days from receipt and conduct a piece count to verify that the component serial numbers agree with those on the shipper’s documents. |  |
| (2) If the material is shipped in containers with tamper indicating seals, the SNMC will verify that the seals are intact and have the correct seal numbers. |  |
| (3) Any differences identified must be reconciled in accordance with SPP-5.8. |  |
| (4) The SNMC reports to Nuclear Fuel the receipt, by serial number or other special identification, of each SNM item. |  |
| 4. Storage, Accountability, and Use of Radioactive Materials | TP-15 |
| a. All licensed radioactive sources and containers of radioactive materials should be properly labeled in accordance with site procedures. Minimum labeling will usually include: |  |
| (1) Individual licensed radioactive source labels will normally contain the standard radiation symbol and warning, an ID number, the isotope and activity, and the activity certification date. |  |
| (2) Containers, such as drums, lockers, cabinets, etc.) used to store licensed radioactive sources and materials will normally have a radioactive material label that includes a general description of the radiological hazards associated with the contents (such as dose rates inside and outside the container, contamination levels, etc.), and any necessary special handling instructions, if appropriate. |  |
| b. All storage locations for BSM sources must be approved by the BSM Controller. These storage locations are established to ensure the physical and administrative control of all licensed BSM. |  |
| c. Administrative controls are established to cover the issue, use, return, and storage of licensed BSM sources. These controls may include: |  |
| (1) Records delegating control of certain sources to a custodian. |  |
| (2) An inventory or list of all sources under the control of the BSM Controller and each Custodian. |  |
| (3) Records documenting the assignment of sources to individual users. | TP-16 |
| (a) Any individual user who signs out a licensed BSM source shall be accountable for the control, handling, and proper return of the source. | HO-05 |

| OE16021, - River Bend Station – A Chemistry technician was using a 20 mCi Xe-133 source in a glass vial on the refuel floor. He had placed the shielded source in the pocket of his dress out clothes. The glass vial clipped out of the lead shield into the pocket of the individual’s protective clothing and was subsequently removed with the protective clothing at the step-off pad. The source was discovered missing when the technician reached the RCA exit area. It was calculated that he received a dose of 42 mR. Error Reduction Tools: Follow Procedures, Self-Checking, Attention to Detail. | Discuss OE16021, Attachment 3 |
| --- | --- |
| (b) An individual user may not transfer assignment of a licensed source to another individual user. The source must be returned to the custodian and assigned to the second user by the BSM Custodian or BSM Controller. | Site procedures may vary. |
| (c) At no time should a non-exempt source be signed out to an individual not in control of and actually using the source or be away from its approved storage location without a user having signed for the source. |  |
| (d) Only individuals authorized by the assigned BSM Custodian or BSM Controller may sign out licensed sources. |  |
| d. Exempt sources may be issued to Custodians and users without the controls placed on licensed sources. | HO-06 Review some of the materials included in this list. |
| (1) The NRC descriptions of exempt concentrations and exempt quantities are outlined in 10 CFR 30.14 and 30.18, respectively. | TP-17 HO-07 & 08, (Attachments 4 & 5) OBJ.-9 |
| (2) The regulations permit, with certain exceptions, exemptions for the receipt, possession, use, transfer, ownership, or acquisition of radioactive products or materials containing byproduct material in concentrations not exceeding those listed in 10 CFR 30.70, or individual quantities each of which does not exceed the applicable quantity set forth in 10 CFR 30.71, Schedule B. |  |
| (3) Only sources received from distributors specifically licensed in accordance with 10 CFR 32 or agreement state requirements to distribute license exempt sources may be considered exempt. | OBJ.-10 |
| e. Control of Special Nuclear Material (SNM). | TP-18 |
| (1) During all activities related to the handling of special nuclear material, the SNM Custodian ensures that RADCON requirements are met. |  |
| (2) All transfers of SNM will be documented. |  |
| (3) Nonfuel SNM will be identified on inventory records. |  |
| (4) Item Control Areas (ICAs) are established for physical and administrative control of SNM. Each ICA has an identifiable physical area such that the SNM being moved into or out of the ICA can be counted and identified. |  |
| (5) The SNM Custodian must authorize any transfers into or out of, or within an ICA. |  |
| (6) If calibration sources containing SNM are removed from an ICA, then only a sign-in/sign-out log is required in lieu of transfer documents. |  |
| (7) Any individual who signs out a SNM calibration source must maintain personal custody of the device until it is returned to the ICA. |  |
| 5. Periodic Monitoring and Inventory of Licensed Radioactive Sources. |  |
| a. As specified in the facility licenses, physical inventories of all licensed BSM sources must be conducted semiannually. The inventories are conducted by the BSM Controller, BSM Custodian, or their designee. |  |
| (1) The inventory is required for all BSM sources containing radioactive material in excess of either 100 µCi of beta and/or gamma emitting material, or 5 µCi of alpha emitting material. | TP-19 OBJ.-11 |
| LER 254-94009, - Quad Cities I – Four sources were not properly controlled. They were removed from the inventory list for disposal, but were subsequently placed back in storage and eventually returned to use. Because they were not on the inventory list, the sources did not receive the required leak test. Causal factors included human performance. Error reduction tools; Self-checking, Attention to detail. | Review LER 254-94009, Attachment 6 |
| (2) Certain radionuclides (e.g., Cs-131, Cr-51, F-18, and Zn-69) have exempt quantities defined in 10 CFR 30, Schedule B, which exceed 100 µCi. Quantities of BSM, in the form of sealed sources, which exceed 100 µCi, must be controlled as licensed and are subject to inventory and leak test requirements. | (See HO-08), Obj. 11 & Obj. 13 |
| (3) Some installed BSM sources cannot be verified in place by direct observation or by detection with survey instrument. For these sources, the BSM Controller will take appropriate action, within the constraints of sound health physics practices and ALARA considerations to verify that the sources are in place. | (Specific site procedures may vary.) |
| (4) If a licensed BSM source is found and/or not documented on the inventory list, the BSM Custodian will: |  |
| (a) Take positive control of the source. |  |
| (b) Immediately notify the BSM Controller. |  |
| (5) When notified that a licensed BSM source is found and/or not documented on the inventory list, the BSM Controller will: | TP-20 |
| (a) Report the deficiency to the RADCON Superintendent. |  |
| (b) Immediately report the deficiency to the Shift Manager who will determine the reportability/reporting requirements. |  |
| (c) Document and report with the assistance of the applicable BSM Custodian, the occurrence and its impact using appropriate corrective action documentation. |  |
| (d) Forward the report to the RADCON Superintendent. |  |
| (6) If a BSM licensed source containing a quantity of radioactivity greater than 10 times the quantity specified in 10 CFR 20, Appendix C, or an aggregate quantity equal to or greater than 1,000 times that quantity, is identified as lost, stolen, or missing and circumstances indicate that an exposure could result to persons in unrestricted areas, the BSM Controller will: | TP-21 OBJ.-12 |
| (a) Document the occurrence and immediately notify the RADCON Superintendent. |  |
| (b) Immediately report the occurrence to the Shift Manager who will determine reportability/reporting requirements. |  |
| For lesser quantities of radioactive material, loss, theft, or missing radioactive material shall be documented and immediately reported to the RADCON Superintendent. |  |
| b. All BSM in use sealed sources containing radioactive material in excess of either 100 µCi of beta and/or gamma emitting material, or 5 µCi of alpha emitting material must be leak tested semiannually. | TP-22 OBJ.-13 |
| (1) Leak tests are also required for certain exempt quantity isotopes as outlined in section 5.a.(2) above. |  |
| (2) Sealed sources that are continuously enclosed within a housing mechanism, such as a boron analyzer or a radiation monitor, need not be tested unless that housing is opened and the source has not been tested within the past six months. |  |
| (3) Sources in storage and not in use are not required to be lead tested as long as they are in storage. However, if the source is placed in use, transferred to another custodian or licensee, or made accessible to personnel, then a leak test must be performed if it has been 6 months or more since the last leak test was performed. |  |
| (4) Start-up sources, excore detectors, and fission detectors not in use must be leak tested within 31 days prior to being installed in their final configuration. |  |
| (5) In addition, leak tests are required: |  |
| (a) Upon receipt of a source from another custodian or licensee, when not accompanied by certification of leak testing within the last 6 months. |  |
| (b) Following repair or maintenance on the source. |  |
| (6) Leak tests are normally performed using dry smears or swabs. If a different medium is necessary, coordinate with the BSM Controller. General guidelines include: | TP-23, OBJ.-14 Consult plant procedures for specific instructions for performing the leak tests. |
| (a) Wipe encapsulated sources directly with a disc smear. |  |
| (b) For plated sources or sources covered with a Mylar window, wipe the container with a disc smear. |  |
| (c) For sources that cannot be removed from its source container, wipe the external surfaces of the container with a disc smear. |  |
| (d) For sources that move through a housing, wipe the housing internals when the source is not present. |  |
| (e) Calculate the activity for each source using the following equation: | TP-24, OBJ.-15 |
| Activity (µCi) = (Gross counts) – BKG counts (Detector Eff.)(2.22 E6) where: Gross Counts = Gross count rate, cpm Bkg Counts = Background count rate, cpm Detector Eff. = Detector Efficiency, 0-1.0 2.22 E6 = 2.22 E6 dpm/µCi. | Error Prevention Tools: Use Self-Checking and Peer-Checking to ensure that data are entered and calculations performed correctly. |
| (7) Direct access to some high activity sources could result in exposure to high dose rates. Smears on the surface housing or most likely path of contamination are considered acceptable leak tests for these sources. | Ask for examples. (Shepherd Irradiator, etc.) |
| (8) If the results of the leak tests (in µCi) demonstrate leakage equal to or greater than 0.005 µCi of radioactive material, then the BSM Controller shall take the following actions: |  |
| (a) Immediately withdraw the source or equipment involved from use or control access until source or equipment can be repaired, replaced, or decontaminated. |  |
| (b) Document and report the occurrence to the RADCON Superintendent as soon as possible using appropriate corrective action documentation. |  |
| c. The SNM Custodian at each site ensures that a physical inventory of all SNM under license is performed at intervals not to exceed 12 months. | TP-25 |
| (1) Physical inventories of ICAs containing SNM reportable under 10 CFR 74.13 (Material Balance Reports) shall all be performed at the same time (having the same physical inventory date) and shall be completed within 30 days of the start date of the inventory. |  |
| (2) The method of inventory will vary depending upon situations such as radiation hazards and the need for protection of equipment. All practical efforts will be made to verify the physical presence of individual items. |  |
| (3) When conditions prevent access to SNM, the inventory may be completed by updating the last physical inventory of record and considering transfers into and out of the ICA. |  |
| (4) Type-E metal seals may be used to ensure the integrity of inventories contained in sealed containers and vaults by verification of the identification number and integrity of the installed seal. |  |
| (5) The SNM Custodian transmits the records to Management Services within 60 days of the inventory. |  |
| d. The SNM Custodian ensures that SNM sealed sources shall be leak tested, as required, before being transferred to another license if the leak test has not been performed within six months. |  |
| 6. Transfer and Disposal of Radioactive Materials | TP-26 |
| a. All non-exempt BSM sources will be returned to the BSM Controller for transfer or disposal. |  |
| (1) The BSM Controller will coordinate with Radwaste when transferring a BSM source to another licensee. |  |
| (2) The BSM Controller will obtain the concurrence of Radwaste prior to disposing a BSM source in a radwaste package. |  |
| (3) Shipments of radioactive materials to other licensees or to disposal sites are normally handled by Radwaste. Shipments are prepared in accordance with the Radioactive Material Shipment Manual (RMSM). | OBJ.-16 |
| b. The SNM Custodian ensures that packages containing SNM for shipment are prepared and controlled in accordance with the RMSM. |  |

| 7. Records | TP-27 |
| --- | --- |
| a. The following records are dispositioned as QA records. They are controlled and retained in accordance with the requirements of SPP-5.9, “Radiological Control and Radioactive Material Shipment Augmented Quality Assurance Program.” |  |
| (1) Original copies of BSM source receipt records. These consist of vendor source certificates and documentation of source type, quantities, and source serial/identification numbers. |  |
| (2) Facility-specific forms for BSM inventories. |  |
| (3) Forms for receipt, offsite transfer, or disposal of BSM sources are handled in accordance with the requirements of the Radioactive Material Shipment Manual. |  |
| (4) Leak test documentation. |  |
| b. Quality-related SNM records are to be retained as long as the licensee retains possession of the material and for 3 years following transfer of the material. |  |
|  | TP-28, Summary |

XII.	SUMMARY

A number of NRC regulations address the control and handling of various forms of radioactive materials, including byproduct material, source material, and special nuclear material.  In addition, TVAN procedures describe the requirements for ensuring that the materials are handled in a safe manner and in accordance with the regulations and the licenses.  For the nuclear power plants, the facility operating licenses, which are a part of the respective plant Technical Specifications, include the authorization for possessing and using Byproduct and Source Materials (BSM) and Special Nuclear Materials (SNM) at the sites.

The RADCON organization at each site is responsible for the control of BSM at the respective facility.  The specific responsibility for directly controlling BSM is assigned to the site BSM Controller.  The BSM Controller is responsible for the receipt, use, and disposition of BSM.  The BSM Controller in turn may assign specific sources to different BSM Custodians within various user groups within the plant.  Individual users may sign out BSM sources from either the BSM Controller or the BSM Custodian within his/her group.

Special Nuclear Material (SNM) at each is controlled by the SNM Custodian from the Reactor Engineering organization.  The SNM Custodian works in coordination with Nuclear Fuels in the Corporate office.  The SNM Custodian is also responsible for inventorying the SNM.

The BSM licenses require the conduct of a routine inventory and leak test of licensed sources.  These inventories and leak tests are performed under the direction of the BSM Controller and Custodians.  The Radwaste Section is responsible for the disposal and offsite shipment of BSM.

Handout # 1

Enabling Objectives

1.	In the context of Radioactive Materials (RAM), define the following:

a.	Byproduct and Source Material (BSM) Controller

b.	Byproduct and Source Material (BSM) Custodian

c.	Byproduct Material

d.	Item Control Area (ICA)

e.	NORM

f.	Physical Inventory

g.	Sealed Source

h.	Source

i.	Exempt Source

j.	Non-Exempt Source

k.	Source Material

l.	Special Nuclear Material (SNM)

m.	User

2.	Identify the document that contains the plant operating license, which also covers the possession and use of BSM and SNM at the site.

3.	Identify the TVAN procedure that addresses the control of BSM.

4.	State the responsibilities of the BSM Controller.

5.	State the responsibilities of the BSM Custodian.

6.	Tell the responsibilities of the SNM Custodian.

7.	Identify the organization that is responsible for receiving materials, including radioactive materials, at each site.

8.	Identify the time period within which packages of radioactive materials received with a Radioactive White I, Yellow II, or Yellow III label must be surveyed.

9.	Identify the NRC regulation that defines exempt concentrations and exempt quantities.

10.	Identify the NRC regulation covering the licensing of persons who allowed to manufacture and distribute exempt quantity sealed sources.

11.	State the requirements for conducting an inventory of BSM.

12.	Tell the actions taken in the event of a lost, stolen, or missing licensed sealed source.

13.	Identify the requirements for conducting leak tests of licensed sealed BSM sources.

14.	Sate the general procedures used in performing leak tests of licensed sealed BSM sources.

15.	Calculate the activity on a leak test smear and identify the activity level on a smear that would indicate that a sealed source is leaking.

16.	Identify the site organization responsible for the disposal of unwanted BSM.

Handout # 2

Terms And Definitions

1.	Byproduct and Source Material (BSM) Controller – That individual within the site RADCON organization or facility RSO organization that performs the BSM source control process.

2.	BSM Custodian – The individual who receives custody of BSM sources for control and use in approved facility activities.

3.	Byproduct Material – Any radioactive material (except SNM) yielded in, or made radioactive by exposure to the radiation incident to, the process of producing or using SNM.

4.	Item Control Area (ICA) – A defined area within a licensed facility with physical boundary delineation wherein Special Nuclear Material (SNM) is kept, subject to material control and accounting practices such that, at any time, an item count and related quantities of SNM can be obtained from the records for the SNM located within the area.

5.	Naturally Occurring Radioactive Material (NORM) – Radioactive material which occurs in nature and is not the product of nuclear fission or a nuclear accelerator.

6.	Physical Inventory – A piece by piece accounting of BSM sources supported by formal documentation.  Acceptable methods of performing the inventory consist of either direct observation of the source or verification by instrument response or other indirect means.  Consolidated sources maintained in sealed containers may be inventoried by ensuring the container is in its assigned location and the seal is unbroken.

7.	Sealed Source – Radioactive material that is encased in a capsule designed to prevent leakage or escape of the material under normal conditions of use and leak tested per facility licenses.

8.	Source, General – Any BSM with known isotopes and activity manufactured for the purpose of measuring, checking, calibrating, or controlling processes quantitatively or qualitatively. These sources include but are not limited to:

A.	Sources in field monitoring equipment to verify operability.

B.	Sources in radiation monitors to maintain a threshold of sensitivity.

C.	Spent uranium slabs used to determine beta response and correction factors for portable monitoring instrumentation.

D.	Sources used to calibrate and response-check field monitoring equipment (portable and fixed).

E.	Liquid standards and liquids and gases used to calibrate and verify calibration of laboratory counting and analyzing equipment.

F.	Sources used for radiographic operations.

9.	Source, Exempt (Non-Licensed) – Those BSM sources exempted from licensing requirements for the user by 10CFR30, 31, and 40. In general, the requirements of 10CFR19, 20, 30, 31, 32, 33, 34, 35, 39, and 40 do not have to be met unless specifically stated in the applicable section. Only sources received from distributors specifically licensed in accordance with 10 CFR 32 or agreement state requirements to distribute license exempt sources, may be considered exempt. Requirements communicated by the sources manufacturer with the “exempt” material are applicable.

10.	Source, Non-Exempt (Licensed) – Any BSM that is not specifically exempted from licensing requirements by the "EXEMPTIONS" section of 10CFR30, parts 30.11 through 30.20, all of 10CFR31, and the "EXEMPTIONS" section of 10CFR40, parts 40.11 through 40.14. A Specific License is required for BSM which does not meet the above exceptions. In general, the requirements of 10CFR19, 20, 30, 31, 32, 33, 34, 35, 39, and 40 must be met.

11.	Source Material – Uranium or thorium, or any combination thereof, in any physical or chemical form, or ores that contain by weight one-twentieth of one percent (0.05%) or more of uranium, thorium, or any combination thereof.  Source material does not include SNM.

12.	Special Nuclear Material (SNM) – Plutonium, uranium 233, uranium enriched in the isotope 233 or in the isotope 235, and any other material which the Nuclear Regulatory Commission, pursuant to the provisions of Section 51 of the Atomic Energy Act, determines to be SNM or any material artificially enriched by any of the foregoing; but does not include source material.

13.	User – An individual who is issued, controls, and uses BSM in performing approved facility activities.

Handout # 3

RESPONSIBILITIES

1	RADCON MANAGER

A. 	Implements applicable procedures.

B.	Ensures compliance with Technical Specifications

C.	Designates a BSM Controller.

D.	Provides technicians to assist the Controller.

2	BSM CONTROLLER

A. 	Maintains the BSM source control program.

B.	Maintains copies of receipt, transfer, disposal, inventory, and leak test records as applicable in an organized, readily accessible file.

3	CUSTODIAN(S)

A. 	Controls all BSM sources assigned to the Custodian.

B.	Provides and maintains a storage area to which only the Custodian or designees have access.

C.	Keeps their inventory and record of issuance to individual users in an organized, readily accessible file.

4.	SNM CUSTODIAN

A.	Manages and controls all SNM.

B.	Maintains copies of receipt, transfer, disposal, inventory, and leak test records relating 
		to SNM in an organized, readily accessible file.

5	USER

A.	Maintains positive control over all BSM sources assigned to the user.

B.	Uses good RADCON and safety practices when using nonexempt BSM sources.

6	MATERIALS AND PROCUREMENT

A.	Maintains all BSM sources in a separate controlled storage location until the RADCON Controller or designee takes custody of the material.

B.	Contacts the BSM Controller upon receipt of BSM sources and assures he receives a copy of all applicable receipt records.

C.	Assures that RADCON/Radwaste receives all applicable original receipt records for BSM sources.

7	RADCON/RADWASTE/ENVIRONMENTAL CONTROL

A.	Does not accept any BSM sources for disposal or transfer to another licensee without the knowledge and approval of the RADCON Controller.

B.	Assures the RADCON Controller receives copies of all applicable shipping papers for BSM sources shipped to another licensee.

C. 	Transfers original receipt and shipping papers to Records Management.

Handout # 4

RADIOLOGICAL LIMITS FOR THE RECEIPT OF RADIOACTIVE MATERIAL

References G and I

NON-EXCLUSIVE USE SHIPMENTS

| PACKAGE DOSE LIMITS mrem/hr |  |  |
| --- | --- | --- |
| TYPE QUANTITY | CONTACT | 1 METER |
| Limited Quantity | 0.5 | N/A |
| Inst. and Articles | 0.5 & 10 at 4" unpackaged | N/A |
| Empty | 0.5 | N/A |
| Type A | 200 | 10 |
| Type B | 200 | 10 |
| PACKAGE CONTAMINATION LIMITS dpm/100cm2 |  |  |
| SMEARABLE | *2,200 - | *220  |

EXCLUSIVE USE SHIPMENTS

| PACKAGE DOSE LIMITS , contact mrem/hr |  |  |  |
| --- | --- | --- | --- |
| TYPE QUANTITY | OPEN VEHICLE |  | CLOSED VEHICLE |
| Inst. and Articles | 2 |  | 2 |
| LSA | 200 |  | 1000 |
| Type A,B, HRCQ | 200 |  | 1000 |
| PACKAGE CONTAMINATION LIMITS dpm/100cm2 |  |  |  |
| SMEARABLE | *22,000 - |  | *2,200  |
| VEHICLE DOSE LIMITS mrem/hr |  |  |  |
| TYPE QUANTITY | CAB | CONTACT | 2 METER |
| Inst. and Articles | 2 | 2 | 2 |
| LSA | 2 | 200 | 10 |
| Type A,B, HRCQ | 2 | 200 | 10 |
| VEHICLE CONTAMINATION LIMITS dpm/100cm2 |  |  |  |
| EMPTY VEHICLE *2,200 -, *220 , 0.5 mrem/hr contact |  |  |  |

*  Some plants have lower limits for some of these values.

Handout # 5

NONEXEMPT SOURCE USER'S RESPONSIBILITIES

INSTRUCTION SHEET

A.	The source user will sign out and sign in all nonexempt sources they use and will assure the Custodian initials each entry.

B.	The source user will maintain control over all sources signed out to the user.

C.	The source user will use the source in compliance with good RADCON and safety practices.

D.	The source user will take proper care of the source to prevent damage to the source. Any damage to a nonexempt source will be immediately reported to the Custodian.

E.	The source user will contact RADCON each time a nonexempt source is to be taken out of the regulated area and/or the protected area.

F.	The source user will contact the Custodian immediately upon loss or theft of a nonexempt source.

G.	If a work activity involving a nonexempt source is to extend beyond the shift of the assigned user, that assigned user must sign the material back in and the designated user on the next shift must sign the source out. At no time will a source be signed out to a person not in control of and actually using the source, or be away from its approved storage location without a user having signed for the source.

Handout # 6

EXEMPT BSM SOURCES AND DEVICES

The following radioactive sources or source containing devices are exemptA from the requirements for licensing and the requirements of SPP-5.6 and site-specific RCIs (except as noted below) per applicable portions of 10 CFR 30, 31D, and/or 40.

ITEM	ISOTOPEB

1. Smoke Detectors 	Am-241, other

2. Exit Signs	Tritium (H-3)

3. Miniature Light Bulb	Kr-85

4. Dewpointers	Ra-226

5. Explosives Detectors	Ni-63

6. Other Self-luminous products	H-3, Kr-85, Pm-147

7. Other Gas & Aerosol detectors	Am-241, Ni-63, others

8. Electron TubesC	H-3, Co-60, Ni-63, Kr-85,

Cs-137, Pm-147

9. Radiation measuring instruments	Various with internal check sources

10. Spark gap irradiators	Co-60

11. Check sources maintained by RADCON	Various for use in verifying 	instrument response.

12. Hydrogen analyzer ion chamber sourcesE	Th-232

A	This exemption only applies to sources which were originally distributed as license exempt in accordance with 10 CFR 32 or a similar Agreement State provision allowing the distribution of license exempt sources. Sources transferred to a site license from another licensee must be considered licensed material regardless of type, form, isotope, or quantity.

B	Sources must meet the applicable constraints and activity limits specified by 10 CFR 30.

C	Electron tubes include spark gap tubes, power tubes, gas tubes including glow lamps, receiving tubes, microwave tubes, radiation detection tubes, indicator tubes, pickup tubes, and any other completely sealed tube that is designed to conduct or control electrical currents.

D	Sources received under General License must be received and maintained in accordance with applicable provisions of 10 CFR 31.

E	These sources are made from commercially available lantern mantles containing naturally occurring thorium and are not under restrictions of federal law.

NOTE: TRANSFER OR DISPOSAL OF ANY EXEMPT RADIOACTIVE SOURCE OR SOURCE CONTAINING DEVICE SHOULD BE COORDINATED WITH THE BSM CONTROLLER AND/OR RADCON.