## Slide 1
ACADs (08-006) Covered
Keywords
Detectors, gas-filled, scintillation, alpha source, beta source, counter test, counting efficiency, minimum detectable activity, percent slope, counter voltage, source check
Description
Supporting Material
1.1.8.1.3 | 1.1.8.1.7 | 3.2.2.1 | 3.2.2.3 | 3.2.3.8.1 | 3.2.3.8.2 | 3.2.3.16 | 3.2.3.23.11
3.2.3.23.6 | 3.2.3.23.7 | 3.2.3.23.8 | 4.4.1 | 4.4.3 | 4.5.4 | 4.5.7.11 | 5.4.1.5b
5.4.1.5c | 5.4.1.5d |  |  |  |  |  | 
Notes: <number>

## Slide 2: Enabling Objectives
Identify 2 Types of Detectors
List 3 Types of Gas Filled Detectors
Name 2 Scintillation Detector Materials
Identify Beta and Alpha Sources Used to Check Laboratory Counters

## Slide 3: Enabling Objectives
State Maximum Voltages for Detectors
Name the Statistical Test Used for Laboratory Counters
Identify the Equation for this Test
State Expected Value for χ2

## Slide 4: Enabling Objectives
State Counting Time for Background
Define Counting Efficiency
Define Minimum Detectable Activity
Identify Conditions that Indicate the Presence of Radioactivity

## Slide 5: Detectors Used in Laboratory Counters
Gas-Filled
Scintillation

## Slide 6: Gas-Filled Detectors
Ionization Chambers
Gas Proportional Detectors
Geiger-Mueller (G-M) Detectors

## Slide 7: G-M Detector with Ludlum 2200 Counter

## Slide 8: Scintillation Detectors
Gamma Detector: Sodium Iodide (Thallium) - NaI(Tl)
Alpha Detector: Zinc Sulfide - ZnS

## Slide 9: ZnS Detector with Ludlum 2200 Counter

## Slide 10: System Reliability Check Sources
Beta – Tc-99
Alpha – Th-230
Traceable to National Institute of Standards and Technology (NIST)

## Slide 11: Characteristics of Tc-99
Long Half-Life (211,000 Years)
Beta Energy: Average: 0.085 MeV Maximum: 0.295 MeV
Results in Conservative Efficiency
Decays by Beta only (No Gamma)
No Daughter In-Growth

## Slide 12: Characteristics of Th-230
Long Half-Life (75,400 Years)
Alpha Particle Energies: 4.68 MeV, 76% 4.62 MeV, 24%
Yields an Efficiency Appropriate with Alphas in the Plant

## Slide 13: ERROR PREVENTION TOOLS
ALWAYS FOLLOW PROCEDURES WHEN HANDLING SOURCES OF RADIOACTIVE MATERIALS!

## Slide 14: Setting the Operating Voltage
CAUTION: Adjust Detector Voltage Slowly. Rapid High Voltage Transits Can Damage Detectors.

## Slide 15: FOCUS ON THE FOUR
Industrial Safety
Equipment Reliability (Take Care of Your Equipment)
Excellence in Human Performance
Intolerance for Equipment Deficiencies (Report Problems)

## Slide 16: Voltage Plateau
Select Lowest Voltage Setting Producing Counts
Insert Source into Counter
Acquire 1-Minute Count
Increase Voltage 50 Volts & Recount
Continue until Counts Rise Dramatically
Plot Data
Select Operating Voltage

## Slide 17: Excessive Voltage
CAUTION Excessive Applied Voltage May Result in Detector Damage. Do Not Exceed the Following Values: Gas Proportional Counters – 1900 Volts G-M Detectors – 1600 Volts Scintillation Counters – 1500 Volts

## Slide 18: Calculate Percent Slope
% Slope = (2 X 104)(C2 – C1) (C1 + C2)(V2 – V1) Where: C1 = cpm at Applied Voltage V1 C2 = cpm at Applied Voltage V2 V1 = Voltage at Base of Plateau V2 = Voltage at Top of Plateau

## Slide 19: Background Determination
Performed Daily
Ensure Operating Voltage Established and Set
Insert Empty, Clean Planchet into Chamber
Acquire 10-Minute Count

## Slide 20: Chi-Square (χ2) Test
Ensure Operating Voltage Established and Set
Ensure Background Determined
Set up Instrument
Acquire Approximately 20 1-Minute Counts (xi)

## Slide 21: χ2 Calculation
Compute Average of 1-Minute Counts (x-bar)
Square Each Count Value (xi)2
Calculate xi – (x-bar)
Calculate [xi – (x-bar)]2
Sum all xi Values

## Slide 22: χ2 Calculation (Continued)
Sum all [xi – (x-bar)]2 Values
Calculate: χ2 = ∑[xi – (x-bar)]2/x-bar
Verify χ2 Within Limits

## Slide 23: Source Check Limits
Performed in Conjunction with χ2 Test
Use Handout 4 Worksheet
Record Lower & Upper Source Check Limits

## Slide 24: Counting Efficiency (Absolute)
Efficiency = (Number of Pulses Recorded)/ (Number of Radiation Quanta Emitted)
Acquire 10-Minute Counts
Calculate Efficiency

## Slide 25: Minimum Detectable Activity
That Activity Which Increases Instrument Readings by an Amount Equal to 3 Times the Standard Deviation of the Background. (Handout 5 Worksheet)

## Slide 26: Summary
Detectors
Source Characteristics
Voltage Plateau
Chi-Square Test
Source Check Limits
Efficiency
Minimum Detectable Activity