## Slide 1
ACADs (08-006) Covered
Keywords
Survey instruments, ionization, gas filled, scintillation, thermoluminesence, geiger-mueller detectors, beta, gamma, neutron, alpha, ion chamber, eberline instruments, frisking device, air sampler, radiation levels, proportional counter, loose surface.
Description
Supporting Material
1.1.8 | 1.1.8.1 | 1.1.8.1.2 | 1.1.8.1.3 | 1.1.8.1.4 | 1.1.8.1.5 | 1.1.8.1.7 | 3.2.2.2
3.2.3.28.1 | 3.3.3.1 | 4.11.1 |  |  |  |  | 
Notes: <number>

## Slide 2: NET 130: Radiological Protection
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
Module 6:
Portable Radiation Survey Instruments

## Slide 3: Overview
Monitoring and measuring radiation levels around the plant
  Colorless
  Odorless
  Tasteless
Q: How can we measure radiation to determine if the levels are safe?
A: By ionization
  Radiation measurement based almost solely on either direct or indirect ionization
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 4: Overview
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
Review of Ionization
  Passage of nuclear radiation through a substance will result in the removal of electrons from neutral atoms
  Ion pair formed (free electron and residual positively-charged atom)
Direct Ionization
  Ion pairs produced by the direct action of charged particles, such as alphas and betas
  Charged particles may collide with electrons and remove them
  If energy transfer insufficient to remove the electron, atom may be left in a state of excitation
Indirect Ionization
  Production of ion pairs by photons, gammas, and X-rays

## Slide 5: Overview
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
Three basic principles of radiation detection:
  Gas-Filled Chambers
  Scintillation
  Thermoluminesence
We will focus primarily on the first two

## Slide 6: Gas-Filled Ion Chambers
Oldest and most widely used device for detecting radiation
Metal cylindrical gas-filled chamber
Voltage (V) applied across two electrodes
  Anode (positive): central electrode
  Cathode (negative): chamber walls
Radiation passing through the chamber ionizes some of the gas atoms
  Positive ions drawn to chamber walls
  Faster-moving electrons (negative) drawn to the central anode
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 7: Gas-Filled Ion Chambers
Electrical charge collects on anode
Charge buildup causes voltage change, aka pulse
Pulse causes current to flow
Ammeter used to measure current
Current level correlated to radiation level
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 8: Gas-Filled Ion Chambers
Pulse magnitude determined by
  Applied voltage
  Number of initial (primary) ionizing events
    Constant number based on specific ionization and detector size
All gas-ionizing types of radiation detectors are technically ion chambers
However, the term “ion chamber” is typically used to refer to such devices for which the input voltage is low enough to prevent secondary ionization (gas amplification)
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 9
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
Six-Region Gas Amplification Curve
, ,  each produce the same detector response.
 _____
- _____
 _____
Voltage
Pulse Height
100
1013
Recombination Region
Ionization Region
Proportional Region
Limited Proportional Region
Geiger-Mueller Region
Continuous Discharge Region
1 2 3 4 5 6

## Slide 10: Gas Amplification
Pulse magnitude determined by
  Applied voltage
  Number of initial (primary) ionizing events
    Constant number based on specific ionization and detector size
  Gas amplification factor for the particular gas used
    Ratio of total # of collected electrons to # of primary electrons
Gas amplification
  After initial (primary) ion pair is formed and moves toward electrodes, secondary ionization occurs
  Additional ion pairs formed
  As applied voltage increases, gas amplification factor changes
The variable that primarily determines the pulse magnitude is the applied voltage
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 11: Geiger-Mueller Detectors
Gas-filled chambers that operate in Region V of the gas amplification curve
Used to detect both beta and gamma radiation
Mylar-covered "window" serves as inlet filter
  Allows passage of  particles of ~0.05 MeV
  Some have movable cover that can be adjusted to allow both  and  to enter, or  only
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 12: Proportional Counter for Neutron Detection
Gas-filled proportional counter detector
Pulse height discriminator
  Distinguish between larger pulses due to neutron radiation (thus secondary ) and smaller pulses due to other incident radiation
Gas used is boron trifluoride (BF3 ) enriched with Boron-10 (10B)
  Very high cross-section for absorbing thermal neutrons
Incident neutrons absorbed by boron atoms
  Results in  emission
   particles cause measureable ionization in the gas
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 13: Scintillation Detectors
Ionizing radiation producing flashes of light when interacting with certain materials containing phosphors
Extremely sensitive to most forms of radiation
  Useful for detecting alpha radiation
  Alpha particles can be detected since they aren’t stopped by the metal walls of a gas chamber
Measures both amount and energy level of radiation present
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 14: Use of Survey Instruments
Portable radiation survey instruments are to be used for the detection and determination of radiation exposure levels.
Durable and able to withstand normal use, but…
Not destruction proof
  Must be handled with care and used properly
  Subject to damage from rough handling the same as any electronic equipment.
By following a few rules and recommendations, the instruments should last many years and provide the user with a measure of knowledge.
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 15: Use of Survey Instruments
Selecting the right instrument for the right job
  Instrument must be capable of measuring the type and intensity of the radiation which is present or suspected.
Beta and gamma radiation
  Ion chamber instruments (Ionization region)
Alpha radiation
  Scintillation detector
  Gas proportional counter (Proportional region)
Neutron radiation
  BF3 gas proportional counter (Proportional region)
Gamma radiation
  G-M detector (G-M region)
  Ion chamber (Ionization region)
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 16: Alpha Survey Instruments
Eberline E-600
Survey instrument commonly paired with scintillator detector
Used to check smears and/or air sample filters for  contamination
Not normally used on daily basis
  Place smear or air sample filter in planchet (specimen cup)
  Place planchet in counting jig
  Place scintillator on counting jig
  Begin test
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 17: Alpha Survey: Eberline E-600
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
Shown with Eberline SHP 380 AB frisker/survey scintillator probe

## Slide 18: Alpha Survey Instruments
Eberline ESP-1
Micro-computer-based survey instrument
Can be set up with different detectors to monitor several types of radiation
For alpha, digital display reads in dpm
Commonly used with AC-3 scintillator probe for alpha detection
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 19: Portable Neutron Survey Instruments
Eberline ASP-2/2E Portable Neutron Rate Meter
Detection and measurement of dose rate (mrem/hr) from both fast and thermal neutron radiation
Measures from 1 to 60000 mrem / hr
Detector: 9” diameter, cadmium-loaded polyethylene sphere with BF3 tube in center
  Allows excellent gamma rejection
  Commonly used: NRD-4 RemBall
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 20: ASP-2/2E with NRD-4 RemBall
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
(different meter shown with ball)

## Slide 21: Portable Neutron Survey Instruments
Eberline ASP-1 with NRD-4
Very similar to ASP-2/2E with NRD-4
Reads 0 to 100000 mrem / hr
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 22: Portable - Survey Instruments
Major portion of radiation encountered in nuclear plant working areas is beta and gamma
Thus, - detecting instruments are most common
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
Teletector Model 6112
Lightweight, portable, battery-operated, beta-gamma survey instrument
Telescoping probe that extends 13 ft to allow dose rate measurements far from source
Measure dose rate due to gamma over wide range of intensities
To detect beta, remove rubber tip from end of probe

## Slide 23: Portable - Survey Instruments
2000W Extender
Similar to Teletector
Extends to 12ft
Audible speaker and illuminated scale
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 24: 2000W Extender
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 25: Portable - Survey Instruments
Rotem Telepole
Similar to Teletector
Extends to 11ft
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 26: Portable - Survey Instruments
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 27: Portable - Survey Instruments
Rotem AMP-100 (Area Monitor Probe)
GM detector with microprocessor-based instrument
Very stable and accurate
Lightweight and compact
Capable of measuring both dose and dose rate for 
Range: 0 to 1000 Rem/hr
Can be used in 3 ways:
  Locally as a hand-held monitor
  Connected to an area monitor for general area dose rates
  Used remotely by means of a wireless remote monitoring system
Can be used for underwater applications, but not typically
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 28: Portable - Survey Instruments
Eberline RM-14 Rad. Monitor
Used most for personnel monitoring
Small, versatile count rate meter
Rechargeable battery or 110 V AC
Adjustable audible alarm
Three multiplier ranges
  500, 5K, and 50K CPM full-scale
  Controlled by 3-position switch:
  Xl, X10 or X100
Response setting:
Fast (~2 sec) or Slow (~20 sec)
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 29: RM-14 Radiation Monitor
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 30: Portable - Survey Instruments
HP-210 G-M tube probe
Commonly used with RM-14
Thin mica window to allow  sensitivity down to ~40KeV
Shielded to permit 
monitoring in a  field
Used for
  Personnel frisker
  Monitoring of table, floors,
  equipment surfaces, etc.
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 31: Portable - Survey Instruments
Eberline RM-14S and RM-14S-A
Newer versions of RM-14
Several scales that earlier model lacked: 50, 500, 5000, 50000, 500000, and 5000000 cpm
All RM-14’s are operated with conservative assumption of 10% efficiency
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 32: Portable - Survey Instruments
Rotem RAM SURF frisking device
Portable, self-contained
Detector and monitor in one
unit
Digital readout with
automatic range adjust
Audible indicator
Also assume 10%
efficiency
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 33: Portable Ion Chamber Instruments
RO-2 / RO-2A / RO-20
Portable air ion-chamber instruments
Chamber vented to atmospheric pressure
Sensitive to , , and X but calibrated for 
Response time: ~5 sec
Single switch turns on instrument, checks batteries, checks zero setting, selects range
Range
  RO-2: 5, 50, 500, and 5000 mr/hr
  RO-2A: 50 mr/hr, 500 mr/hr, 5 R/hr, and 50 R/hr
  RO-20: all of the above
For : open shield to expose mylar window
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 34: RO-2 RO-20
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 35: Portable Ion Chamber Instruments
Bicron RSO-5 / RSO-50
Similar to RO-2 and RO-2A
  RSO-5: Same scales as RO-2
  RSO-50: Same scales as RO-2A
Measuring  and  dose rates (includes  window)
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 36: Portable Ion Chamber Instruments
Rotem RAM ION
Similar in function to RO-2 or RSO-5
Case similar to RAM SURF
Portable, self-contained
Detector and monitor in one unit
Digital readout with automatic range adjust
Range 0.1 mRem/hr to 50 Rem/hr
Beta window
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 37: Portable Ion Chamber Instruments
Rotem RAM ION
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 38: Other Devices
RADēCO H809V Air Sampler
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 39: Other Devices
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 40: Portable Instruments for Monitoring For Loose-Surface Contamination
For beta and gamma: RM-14, RM-14s, and RAM SURF are used for determination of quantity
For alpha: E-600 and ESP-1
E-600 probe must be held within 1/4 inch of the area being monitored to ensure proper response.
RM-14 with HP-210 probe
  Used for monitoring personel contamination
  Probe must be within 1/2 inch of the surface being monitored to pick up possible contamination.
E-600 and RM-14 with HP-210 require slow probe movement to cover the area being monitored, about 1 to 2 in / sec
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 41: Meter Reading Challenge
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 42: CPM ?
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 43
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 44: CPM and DPM
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 45
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 46
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 47
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 48
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 49
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 50: CPM and DPM and uCi Efficiency=18%
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 51: Meter Reading Challenge
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
Let’s compare notes…..

## Slide 52: CPM ?
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 53: # 1= 3200 cpm
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 54
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 55: 10,000 cpm
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 56: CPM and DPM
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 57: 22,000 cpm
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
DPM=cpm/efficiency
22,000 cpm/.10
220,000 dpm

## Slide 58
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 59: 300 cpm
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 60
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
mR/hr

## Slide 61: 400 mR/hr
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 62
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 63: 3.6 mR/hr
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 64
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 65: 2.5 R/hr
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 66
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 67: 8.5 R/hr
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 68: CPM and DPM and uCi Efficiency=18%
NET 130
Module 6: Portable Radiation Survey Instruments
<number>

## Slide 69: 25.1 cpm
NET 130
Module 6: Portable Radiation Survey Instruments
<number>
DPM=CPM/Efficiency
=25.1/0.180
=139 DPM