# NET 130:
Radiological Protection

## Module 5:

## Radiation Detection Principles and Instruments

# Overview

## Many instruments have been developed to detect radiation

## Based on knowledge of how radiation interacts with matter

### Excitation

### Ionization

#### Charged particles cause ionization directly through Coulombic interactions

#### EM radiation produces ion pairs in matter

##### Photoelectric effect

##### Compton scattering

##### Pair production

#### Neutrons produce ions through secondary mechanisms

## Four methods for detecting ionizing radiation:

### Ions collected to produce signal

### Amplification of ionization to produce stronger signal

### Fluorescence of a substance that has absorbed energy from radiation

### Radiation-induced chemical reactions

## Three major types of detection instruments:

### Nuclear instrumentation

### Portable survey instruments and area monitors

### Personnel monitoring devices

# Gas-Filled Detectors

## Detect incident radiation by measurement of two ionization processes

### Primary process: ions produced directly by radiation effects

### Secondary process: additional ions produced from or by effects of primary ions

#### Townsend Avalanche

## Primary and secondary ions produced within the gas are separated by Coulombic effects and collected by charged electrodes in the detector

### Anode (positively charged electrode)

#### Collects the negative ions

### Cathode (negatively charged electrode)

#### Collects the positive ions

# Gas-Filled Detector: Components

## Cylindrical gas chamber

### Air

### P-10 gas mixture (10% methane, 90% argon)

### Helium

### Neon

## Anode (+):	Wire at center of chamber

## Cathode (-):	Chamber walls

## Operating Principles

### Voltage applied across electrodes

### Incident radiation (α, β, or γ) enters chamber and ionizes the fill-gas

### Ions (+/-) separate and migrate to respective electrodes

### Current output is generated and scaled to radiation level

## Voltage too low

### Ions may recombine and neutralize each other prior to reaching electrodes

## Proper operating voltage

### All primary ion pairs are collected

## Voltage too high

### Chamber becomes flooded with ions due to secondary ionizations caused by high-energy primary ions

### Output current is no longer proportional to number of primary ionizations

### Radiation events no longer measured

#### Ionization “avalanche” propagated by input voltage itself

## Recombination Region

### Applied voltage too low

### Recombination occurs

### Low electric field strength

## Ionization Chamber Region

## (aka Saturation Region)

### Voltage high enough to prevent recombination

#### All primary ion pairs collected on electrodes

### Voltage low enough to prevent secondary ionizations

### Voltage in this range called saturation voltage

### As voltage increases while incident radiation level remains constant, output current remains constant (saturation current)

## Proportional Region

### Gas amplification (or multiplication) occurs

#### Increased voltage increases primary ion energy levels

#### Secondary ionizations occur

#### Add to total collected charge on electrodes

### Increased output current is related to # of primary ionizations via the proportionality constant

### (aka gas multiplication factor)

#### Function of detector geometry, fill-gas properties, and radiation properties

## Limited Proportional Region

## Collected charge becomes independent of # of primary ionizations

## Secondary ionization progresses to photoionization (photoelectric effect)

## Proportionality constant no longer accurate

## Not very useful range for radiation detection

## Geiger-Mueller (GM) Region

### Any radiation event strong enough to produce primary ions results in complete ionization of gas

### After an initial ionizing event, detector is left insensitive for a period of time (dead time)

#### Freed primary negative ions (mostly electrons) reach anode faster than heavy positive ions can reach cathode

#### Photoionization causes the anode to be completely surrounded by cloud of secondary positive ions

#### Cloud “shields” anode so that no secondary negative ions can be collected

#### Detector is effectively "shut off"

#### Detector recovers after positive ions migrate to cathode

### Dead time limits the number of radiation events that can be detected

#### Usually 100 to 500 s

## Continuous Discharge Region

### Electric field strength so intense that no initial radiation event is required to completely ionize the gas

### Electric field itself propagates secondary ionization

### Complete avalanching occurs

### No practical detection of radiation is possible.

# Gas-Filled Detectors

## Most commonly used detection instrument due to versatility

### Can detect and discern between all types of radiation over entire energy spectrum

### Cylindrical shape provides the strongest electric field and output current for a given operating voltage

## Most common detectors operate in the ionization chamber, proportional, and Geiger-Mueller regions

## No detectors operate solely in the recombination, limited proportional, or continuous discharge regions.

## Can discriminate between , , and  radiation

### Pulse height discrimination: electronically filter out pulses below or above expected height for radiation type of interest

## Less sensitive over long range than GM

## Include:

### Portable neutron radiation survey meters

### Personnel contamination monitoring

## Include:

### Area radiation monitors

### Portable high-range radiation survey meters (Teletector)

## Advantages

## highly sensitive: capable of detecting low intensity radiation fields

## Only simple electronic amplification of the detector signal is required

## less insulation required to decrease “noise” interference

## Some GM detectors detect  only

### Solid casing

## Some detect , and 

### ,  radiation: short travel range

#### Cannot penetrate detector casing

### Mylar window to allow  and  radiation to enter

###  and  can be separately detected by using different window types and thicknesses to filter incident radiation

### Shield must be placed over window to detect 

#### Blocks  and 

# Scintillation Detectors

## Detect radiation by induction of luminescence

### Absorption of energy by a substance with the subsequent emission of visible radiation (photons)

## Incident radiation interacts with the scintillator material

## Excites electrons in material

## Electromagnetic radiation emitted in the visible light range

## Common scintillator materials

### Anthracene crystals

### Sodium iodide crystals

### Lithium iodide crystals

### Zinc sulfide powder

### Lithium iodide, boron, and cadmium

### can be used to detect neutrons

# 6 Steps of Scintillation Detection

## Inside scintillator:

### Excitation due to absorption of radiation

### Emission of light photons from de-excitation

### Transit of light to photocathode inside photomultiplier tube

## Inside photomultiplier tube:

### Production of photoelectrons in photocathode

### Multiplication of photoelectrons

## Outside scintillator and photomultiplier tube:

### Conversion of electronic detector output to useful information

# Common Scintillator Materials

## Anthracene crystals

## Sodium iodide crystals

## Lithium iodide crystals

## Zinc sulfide powder

## Lithium iodide, boron, and cadmium can be used to detect neutrons

# Photocathode

## Light-sensitive material that absorbs photons and emits photoelectrons

## Common material: Antimony-Cesium

## Emits about one electron for every 10 photons absorbed

# Photomultiplier Tube: Dynodes

## Photoelectrons strike successive dynodes and are multiplied (secondary electron production)

## Amplifies the output signal

## If tube has 10 dynodes, total gain would be around 106

## Typical tubes made with 6 to 14 dynodes

# Semiconductor Detectors

## Operation similar to gas-filled detectors, but chamber filled with solid semiconductor material

## Crystalline material whose electrical conductivity is intermediate between that of a good conductor and a good insulator

## Benefits compared to other types

### Very little fluctuation in output for a given energy of radiation

### Fast

## Energy transfer from radiation to semiconductor target produces a freed electron and an electron vacancy, or hole

## Electrons travel to the anode

## Hole “travels” toward the negative electrode

### Not physically

### Successive exchanges of electrons between neighboring molecules in the crystalline lattice

# Semiconductor Detectors: Pros/Cons

## Pros

### Fast response time

#### Due to high mobility of electrons and holes

#### Takes longer for ions to physically travel through space in a gas-filled detector

### Less statistical fluctuations for any given radiation energy

#### A smaller amount of energy required to produce electron-hole pair in a semiconductor than an ion pair in a gas

#### For a given energy, 8 to 10 times as many charge-carrying pairs are produced in semiconductors as in gases

### Total charge collected varies linearly with radiation energy

## Cons

### Very sensitive to heat: must be cooled to eliminate error

### Photomultiplier output very weak

#### Powerful amplifiers needed in the external circuit

# Detection Systems

## Two main components:

### Detector

#### Gas-filled, scintillation, or semiconductor

### Measuring apparatus

#### Converts signal output from detector to usable information for the operator

## Detection system categories, by output type:

### Pulse-type output

### Mean-level output

## Detection system categories, by application:

### Nuclear instrumentation

### Portable survey instruments and area monitors

### Personal dosimetry

## Pulse-Type Output:

### records a series of individual signals (pulses) separated or “resolved” over time

### each pulse represents a separate radiation event within the detector

### “Frisker”-type survey instruments found near any contaminated area access point

# Nuclear Instrumentation (NI)

## NI detectors are used to measure/record neutron (η) flux as a measure of reactor power level

## Range of η flux is wide, spanning from:

### Shutdown

### Reactor start-up

### 100% power

## To accurately monitor η population at all power levels, there are three overlapping detector ranges

### Source range: (100 – 106)

### Intermediate range: (101 – 1010)

### Power range: (1010 – 1012)

# NI Detector Ranges

# Neutron Energy Ranges

## Fast neutrons have an energy > 1 eV

## Slow neutrons have an energy less than or equal 0.4 eV.

## Hot neutrons have an energy of about 0.2 eV.

## Thermal neutrons have an energy of about 0.025 eV.

## Cold neutrons have an energy from 5x10−5 eV to 0.025 eV.

# NI:  Fission-Chamber Detectors

## Neutron detection in source and intermediate ranges

## Gas-filled ionization-type detector

## Inner “cans” coated with U-235 lining

## Fast neutrons exiting the core are thermalized by the time they make their way inside the F-C detector

### Interact with materials outside the core

### Interact with the plastic covering of the detector

## Thermal neutrons lead to fission of  the U-235 lining inside the detector

## Reactor core neutron flux is then measured as a product of the fission of U-235 in the F-C detector

# Fission-Chamber Output Signal

## Pulse height discrimination implemented in order to pass only the signal portion due to neutron effects

## Pulse discriminator bias: the selective value for pulses

## Products of incident thermal neutrons: fission fragments with average energy of 165 MeV

## Energy of alphas from uranium isotope decay: 4 MeV

## Fission gammas: no more than 7 MeV

## The fission fragment energy due to neutron entering the detector is clearly distinct

### Pulse is much larger than those for non-fission reactions within detector

# Pulse Height Discrimination

# NI: Power vs. Intermediate Range

## Any power level: reactor produces both neutron and gamma fluxes

## In intermediate range, exact correlation between gamma and neutron flux is not easily predictable

## For an ion chamber to read power in the intermediate range, it must be compensated

### Electronically cancel out gamma effects

## In power range, gamma flux becomes insignificant compared to neutron flux

## Gamma compensation no longer necessary

# NI: Uncompensated Ion Chambers

## Monitor reactor power in the power range

### Single boron-lined cylindrical chamber operating in the ionization chamber region

### Mean-level output

### Gamma-induced current typically represents only 1% of total output signal

# NI: Incore Instrumentation

## Monitor power production at select locations within the core

## Verify reactor core design parameters: flux mapping

## Data only– no operational plant control

## Simpler version of fission chamber

### Approx 0.2” diameter, 2.1” length

### Uses uranium oxide clad in stainless steel, with helium fill gas

# NI Detector Circuitry

## A channel consists of a detector, its measuring apparatus (transducer), and a display

### Sends signals to the reactor control and protective systems

## At main control panel, reactor core is monitored by

### Two source range channels

### Two intermediate range channels

### Four power range channels (0 to 120% power)

## Third source range channel with dual displays

### Nuclear instrument cabinet

### Control room evacuation panel

## Main control panel: U-235-based FC detector

## Instrument cabinet: Boron Triflouride (BF3) FC detector

## Both are dual-element (dual-can) detectors

### Provides increased sensitivity in the low  fluxes of the source range

## Pulse height discriminators “screen out”  flux from 

## Each FC is powered by a high voltage power supply

## Each FC output is amplified and filtered by a separate preamplifier

### Filter electronic noise due to cable lengths

## Preamplifier outputs from 2 FCs are summed at the channel’s discriminator

### Non-neutron pulses are filtered out

### Signal is further processed and amplified for use as indication of power level

# Nuclear Instrumentation: Intermediate Range

## Spans the source (100 – 106) and power (1010 – 1012) ranges

## One U-235 fission chamber (can) per detector

## Output signal

### Pulse-type in the source range

### Mean-level in the power range

# Nuclear Instrumentation: Intermediate Range

## “Pulse pile-up”

### When passing source and power ranges in either direction, neutron events occur and change so rapidly that less overall sensitivity is needed

### Gamma flux is not predictably related to neutron flux

### Cannot be “filtered” by pulse height discrimination

### Intermediate range neutron flux levels are several orders of magnitude higher than source

### Pile-up occurs at upper end of detector range due to high magnitude of combined  and  flux

### The predictable pulses from the lower end effectively change from an AC signal to a fluctuating DC signal

## Campbell Theorem

### “With a random occurrence, the variations in the occurrence is proportional to the square root of the random rate."

### Simply put: by taking the mean value of the oscillations in detector output appearing at the upper-end of the detector scale, a meaningful detector signal is obtained

## Monitored by four independent channels

## Each channel uses a long, boron-lined uncompensated ion chamber

## Each chamber includes two separate neutron detecting sections

## Gammas are so out-numbered by the neutrons that gamma-compensation is not necessary

## One high voltage supply (0-1500 VDC per channel) powers both sections of detector

## Output current from each section is fed to an amplifier

## Amplifier output sent to

### Protection and control systems

### Control panel readouts

### Summing amplifier

#### Add signals from separate detector sections and amplify to make combined signal proportional to total core power (0 to 120%)

## Summing amplifier output sent to

### Protection and control systems

### Control panel readouts

## Gammas are so out-numbered by the neutrons that gamma-compensation is not necessary

# Power Range Detector Channel

# Portable Survey Instruments
and Area Monitors

## Survey meters

### Compact detector systems used to monitor an area for neutrons, beta, alpha, or gamma radiation

## Portable Instruments

### Survey meter, powered by batteries

### Can be carried to any remote location

## Area Monitors

### Survey instruments in permanent installation

# Portable Survey Instruments
and Area Monitors

## Considerations:

## Reduction of size and weight

### Gas-filled detector produces the most intense output for the lowest applied voltage

### To reduce the size and weight

#### Reduce battery size: balance between weight and battery life

#### Reducing chamber size: erratic readings, useless

### A bulky, reliable instrument is preferred to a small one that yields erratic results

## Type, energy, and intensity of the radiation field

### Low range beta and gamma survey meters

### High range beta and gamma survey meters

### Alpha survey meters

### Neutron survey meters

# Low Range  and  Survey Meters

## Low range: fields ranging from background level to levels of a few hundred milliroentgens per hour

## Most used: Geiger-Mueller tube

## Also used: Scintillation detectors

# Low Range  and  Survey Meters

## G-M tube: Advantages

### Variety of sizes and shapes

### Inexpensive

### The slightest radiation event strong enough to cause primary ionization results in ionization of the entire gas volume

### Thus detector is highly sensitive, even in lowest intensity radiation fields

### Only simple electronic amplification of the detector signal is required

#### Hardware lasts longer

#### Requires less power

### Strong output signal means G-M needs less electrical noise insulation than other detectors

# Low Range  and  Survey Meters

## G-M tube: Disadvantages

### Incapable of discerning between type and energy of the radiation event

### Only counts events and yields output in events per unit time or dose rate

### A beta particle or gamma ray, high or low energy, represents one event counted

### Only capable of detecting fields to some upper limit of intensity

#### Limited to lower intensity fields due to detector dead time

## Most common G-M gases: noble gases

### Helium

### Argon

### Neon

### Sometimes hydrogen and nitrogen

### Characteristics of gas affect dead time

## After primary ionization, avalanche, and output pulse, G-M detector enters phase called tube recovery

### Positive ions slowly migrate to cathode

### Neutralized upon arrival

### Neutralization may result in production of additional electrons and/or photoelectrons

### Can result in another discharge of the tube, effectively lengthening dead time

## Quenching

### Process used to prevent multiple G-M tube discharges

### Methods

### Electronic circuitry external to detector (inefficient)

### Quench gases added to the gas volume

#### Self-quenching, efficient

#### Common type:  ethyl alcohol, bromine or chlorine

#### Quench gas molecules neutralize positive ions in fill-gas before they can reach cathode

#### Charged quench gas molecules are then neutralized by cathode

#### Dampens potential for secondary discharge

## G-M tube requires high input voltage

### Permits strong signal from ion collection

### Frequent replacement of high voltage batteries

## Detecting beta particles with G-M

### Particles have short range: window required

### Mica, mylar, or thin stainless steel

### Based on window material and thickness, correction factors can be determined to help narrow output to reflect beta activity alone

# High Range  and  Survey Meters

## Most: Uncompensated ion chambers

### Very simple compared to G-M

### Pulse or mean-level output

### Strength of output signal is directly proportional to the # of ion pairs collected

### Correlates in turn to a function of radiation energy

### Signal converted to dose rate

## Can detect wide range of field intensities, but…

## Disadvantages of Ion Chambers

### Output signal weak

### Must be amplified considerably

### Battery power limitations

### Electronic noise-- frequent zeroing (taring) might be required

### Signal-to-noise ratio renders ion chamber inefficient at low range compared to G-M

## Ion chambers

### Insulation must be very good

### Fill gas

#### Air at atmospheric pressure

#### Noble gases

### For beta, ion chamber must be equipped with a thin wall or window

# Alpha Survey Instruments

## A 1.0 MeV alpha particle has a range in air of only ~0.6 cm compared to 330 cm range of beta particles of the same energy in air

## Alpha particle ranges are considerably shorter in denser materials

## Detectors commonly used:

### Ion chambers

#### Small field intensities

#### Very thin window must be incorporated

### Scintillation detectors

#### Very effective

## Scintillation detectors (continued)

### Commonly sodium iodide (NaI), cesium iodide (CsI), or silver-activated zinc sulfide also

### Activator materials

#### Desirable “impurities” in scintillator material

#### Capture electrons and holes created through ionization of the scintillator and to emit the light photons upon returning to ground state

#### Examples

##### Silver in zinc sulfide

##### Thallium in sodium and cesium iodides

# Neutron Survey Instruments

## Neutrons alone do not produce ionization (a detectable signal)

## The detector must contain a material with which the neutron interacts to produce ions

## Most common target material: Boron

### Either as a fill gas or a coating on the inner wall

## Survey meters used for neutron detectors

### Gas-filled

### Semiconductor

### Scintillation

## Common setup for gas-filled portable neutron survey meter

### BF3 proportional counter surrounded by a cadmium loaded, polyethylene sphere

#### Sphere thermalizes incoming fast neutrons so they can be detected

### Meter output can read directly in millirem or rem per hour

## Ion chambers as neutron survey instruments

### As with NI, commonly use boron-10 coating

### Must be compensated for gamma

## Scintillation detectors as neutron survey instruments

### Common material: lithium iodide

### Thermal neutrons interact with lithium to form tritium and an alpha particle

### Alpha particle causes measurable ionization

# Portable Radiation Survey Instruments

## Pre-operational checks:

### Battery-check: battery strength should be well within the acceptance range

### Calibration check: a sticker affixed to the side of the instrument notes the “calibration due date”

### Visual inspection: inspect for no visible signs of damage (i.e. loose, or missing parts; damaged detector; cracked meter face, etc.)

### Source check: expose the detector to the “check source” [either internal or external] and note the meter response. It shall respond within the set “acceptance range”

## Pre-operational checks:

### Zero-check: adjust so that the meter reads zero (0) when unexposed to a source

### Light-check: applicable to a  scintillation detector, the meter should read zero (0) when no source is present. Spurious counts are indicative of a damaged detector (i.e. light is leaking into the detector)

# Pocket Ion-Chamber Dosimeter

## Self-reading

## Records dose in either milliroentgen (mR) or Roentgen (R)

## Sensitive to gamma only

## DC voltage applied to quartz fibers inside the chamber

## As gamma interactions/ionizations occur within the chamber, and the ions are collected, the coulombic repulsion is decreased, and the fibers move closer together

## Extremely sensitive to shock

## Located in Emergency Kits

# Electronic Dosimeter

## Activated using the Electronic Dosimeter reader

## Self-reading, digital display of accumulated dose and dose rate

## Records accumulated dose and “highest dose rate field” in milliroentgen (mR) and milliroentgen per hour (mR/hr), respectively (can be set to record in units of Roentgen)

## Visual and audible alarms for accumulated dose and dose rate

## Silicon diode detector to detect gamma radiation (sensitive to gamma only)

## Resistant to shock

## Downloads dose and dose rate data to Radiation Exposure Control database when deactivated by the ED reader

# Thermoluminescent Dosimeter

## Scintillation-type device

### Lithium Flouride (LiF) or Calcium Flouride (CaF)

## CaF: high efficiency for detecting gamma and “X” radiation – poor at detecting neutrons

## LiF:  capable of detecting alpha (α), beta (β), gamma (γ),“X”  and neutron (η) radiation – typically used at nuclear power plants

## As incident radiation interacts with the crystal, the resultant ionization/excitation energy is stored.

## Crystal retains this energy until heat is applied.

## The “trapped” energy is then released in the form of light, as the atoms of the crystal return to their “ground state”

## The light emitted is then correlated to dose received

## Once the TLD has been “read”, memory is cleared

## TLD is then available for re-use